LCLS Injector Quadrupole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Roger Carr |
| Date: | Thursday, June 01, 2006 |
| Work Order/Charge No.: | 92-4215-8 |
| Serial Number | 000391 |

Part Set-up – Coordinate System Set-up

Planar Alignment

- Mid-Plane of the magnet

Spatial Alignment

- A line on the top part of the magnet

  - +X goes towards (Magnet Info Label)

“Z” Zero

- Mid-Plane of the magnet

“X” & “Y” Zero

- On both ends

  - Tangent point of each radii (4 on each end, 8 total).

    - Create a line between diagonal tangent points creates 2 lines.

      - Intersect the lines.

        - Creates a point on each end.

- Create a line of these 2 end points

  - This is the “X” & “Y” Zero, and Beamline or” Z” Axis.

Tooling Ball Measurements/Locations

Top of magnet; view from “+Y”

| Tooling Ball | FORM | DIAMETER | X | Y | Z |
| --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00030 | 0.49908 | 6.50004 | 8.87887 | -1.25164 |
| TB 2 | 0.00006 | 0.49997 | -6.49905 | 8.87804 | -1.25049 |
| TB 3 | 0.00011 | 0.49998 | -6.49786 | 8.87861 | 1.24881 |
| TB 4 | 0.00035 | 0.49785 | 6.50056 | 8.87998 | 1.24962 |

Additional Requested Measurements

Tangent Point Straightness

- 0.00024

- 0.00009

- 0.00037

- 0.00032

Parallelism to Beamline

- 0.00049

- 0.00037

- 0.00069

- 0.00050

Distance of opposite Tangent

Axis

A-C = 1.26050 (0.63023, 0.63027)

B-D = 1.25828 (0.62930, 0.62937)

# Text boxes

TB 1

TB 2

TB 3

TB 4

Z+

+Z

+X

A

B

C

D

View From +Z