LCLS Injector Quadrupole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | T. Borden |
| Date: | Wednesday, January 10, 2007 |
| Work Order/Charge No.: | 92-4215-8 |
| Serial Number: | 002402 |
| URL of Fiducial Report: | \\Web002\www-group\met\Quality\FIDUCIAL REPORTS\LCLS QUADS\LCLS QUAD 002402.pdf |

Part Set-up – Coordinate System Set-up

Planar Alignment

- Mid-Plane of the magnet

Spatial Alignment

- A line on the top part of the magnet

  - +X goes towards (Magnet Info Label)

“Z” Zero

- Mid-Plane of the magnet

“X” & “Y” Zero

- On both ends

  - Tangent point of each radii (4 on each end, 8 total).

    - Create a line between diagonal tangent points creates 2 lines.

      - Intersect the lines.

        - Creates a point on each end.

- Create a line of these 2 end points

  - This is the “X” & “Y” Zero, and Beamline or” Z” Axis.

Tooling Ball Measurements/Locations

Top of magnet; view from “+Y”

| Tooling Ball | FORM | DIAMETER | X | Y | Z |
| --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00108 | .49817 | 6.49805 | 8.87611 | -1.25192 |
| TB 2 | 0.00100 | 0.49735 | -6.50354 | 8.87587 | -1.25231 |
| TB 3 | 0.00061 | 0.49782 | -6.50144 | 8.87579 | 1.24946 |
| TB 4 | 0.00006 | 0.49999 | 6.49734 | 8.87492 | 1.24926 |

Additional Requested Measurements

| Pole | Straightness | Parallelism |
| --- | --- | --- |
| A | 0.00039 | 0.00033 |
| B | 0.00030 | 0.00019 |
| C | 0.00028 | 0.00023 |
| D | 0.00020 | 0.00068 |

| Pole Dist | Z+ Side | Z- Side |
| --- | --- | --- |
| A-C | 1.26060 | 1.26030 |
| B-D | 1.26010 | 1.26007 |

# Text boxes

TB 1

TB 2

TB 3

TB 4

Z+

+Z

+X

A

B

C

D

View From +Z