LCLS Injector Quadrupole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | T. Borden |
| Date: | Thursday, January 11, 2007 |
| Work Order/Charge No.: | 92-4215-8 |
| Serial Number: | 002403 |
| URL of Fiducial Report: | \\Web002\www-group\met\Quality\FIDUCIAL REPORTS\LCLS QUADS\LCLS QUAD 002403.pdf |

Part Set-up – Coordinate System Set-up

Planar Alignment

- Mid-Plane of the magnet

Spatial Alignment

- A line on the top part of the magnet

  - +X goes towards (Magnet Info Label)

“Z” Zero

- Mid-Plane of the magnet

“X” & “Y” Zero

- On both ends

  - Tangent point of each radii (4 on each end, 8 total).

    - Create a line between diagonal tangent points creates 2 lines.

      - Intersect the lines.

        - Creates a point on each end.

- Create a line of these 2 end points

  - This is the “X” & “Y” Zero, and Beamline or” Z” Axis.

Tooling Ball Measurements/Locations

Top of magnet; view from “+Y”

| Tooling Ball | FORM | DIAMETER | X | Y | Z |
| --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00298 | 0.48523 | 6.50082 | 8.87918 | -1.25460 |
| TB 2 | 0.00282 | 0.48863 | -6.49314 | 8.87829 | -1.25444 |
| TB 3 | 0.00047 | 0.49796 | -6.49393 | 8.87506 | 1.24765 |
| TB 4 | 0.00006 | 0.50005 | 6.50033 | 8.87364 | 1.24902 |

Additional Requested Measurements

| Pole | Straightness | Parallelism |
| --- | --- | --- |
| A | 0.00032 | 0.00031 |
| B | 0.00031 | 0.00045 |
| C | 0.00032 | 0.00031 |
| D | 0.00019 | 0.00012 |

| Pole Dist | Z+ Side | Z- Side |
| --- | --- | --- |
| A-C | 1.26106 | 1.26070 |
| B-D | 1.25968 | 1.25918 |

# Text boxes

TB 1

TB 2

TB 3

TB 4

Z+

+Z

+X

A

B

C

D