LCLS Injector Quadrupole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | T. Borden |
| Date: | Thursday, January 11, 2007 |
| Work Order/Charge No.: | 92-4215-8 |
| Serial Number: | 002404 |
| URL of Fiducial Report: | \\Web002\www-group\met\Quality\FIDUCIAL REPORTS\LCLS QUADS\LCLS QUAD 002404.pdf |

Part Set-up – Coordinate System Set-up

Planar Alignment

- Mid-Plane of the magnet

Spatial Alignment

- A line on the top part of the magnet

  - +X goes towards (Magnet Info Label)

“Z” Zero

- Mid-Plane of the magnet

“X” & “Y” Zero

- On both ends

  - Tangent point of each radii (4 on each end, 8 total).

    - Create a line between diagonal tangent points creates 2 lines.

      - Intersect the lines.

        - Creates a point on each end.

- Create a line of these 2 end points

  - This is the “X” & “Y” Zero, and Beamline or” Z” Axis.

Tooling Ball Measurements/Locations

Top of magnet; view from “+Y”

| Tooling Ball | FORM | DIAMETER | X | Y | Z |
| --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00114 | 0.49618 | 6.49823 | 8.87835 | -1.24699 |
| TB 2 | 0.00088 | 0.49584 | -6.50008 | 8.87831 | -1.24507 |
| TB 3 | 0.00131 | 0.49669 | -6.50048 | 8.87835 | 1.25483 |
| TB 4 | 0.00029 | 0.49942 | 6.49904 | 8.87701 | 1.25125 |

Additional Requested Measurements

| Pole | Straightness | Parallelism |
| --- | --- | --- |
| A | 0.00042 | 0.00040 |
| B | 0.00021 | 0.00064 |
| C | 0.00040 | 0.00033 |
| D | 0.00024 | 0.00022 |

| Pole Dist | Z+ Side | Z- Side |
| --- | --- | --- |
| A-C | 1.25986 | 1.25919 |
| B-D | 1.25919 | 1.25970 |

# Text boxes

TB 1

TB 2

TB 3

TB 4

Z+

+Z

+X

A

B

C

D