LCLS Injector Gun Solenoid Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Roger Carr |
| Date: | Friday, May 05, 2006 |
| Work Order/Charge No.: | 92-4215-8 |
| Serial Number | 1 |

Part Set-up – Coordinate System Set-up

Planar Alignment

- Mid-Plane of the magnet

Spatial Alignment

- A line on the TB hole side of the Upstream End Plate measured in the vertical from base.

  - +Y goes towards (Magnet Label Sidel)

“Z” Zero

- Mid-Plane of the magnet

“X” & “Y” Zero

- On both ends

  - Center of each of the End Plate bores.

    - Create an Axis (Beamline) of the two points

  - This is the “X” & “Y” Zero, and Beamline or” Z” Axis, where +Z is towards the opposite the tooling ball end plate.

| 3-D View | Front View |
| --- | --- |

Tooling Ball Measurements/Locations

Upstream Endplate (TB 1-5)

| +X Side (TB 6 & 7) | -X Side (TB 8 & 9) |
| --- | --- |

| Tooling Ball | Form | Diameter | X | Y | Z |
| --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00030 | 0.49834 | -5.79831 | 5.79992 | -5.09461 |
| TB 2 | 0.00025 | 0.49774 | -8.20053 | -1.49955 | -5.09453 |
| TB 3 | 0.00039 | 0.49860 | -5.79932 | -5.79441 | -5.09617 |
| TB 4 | 0.00030 | 0.49860 | 5.79312 | -5.79476 | -5.10309 |
| TB 5 | 0.00033 | 0.49955 | 5.79504 | 5.79767 | -5.09711 |
| TB 6 | 0.00028 | 0.49847 | 6.85796 | 6.86143 | -3.21855 |
| TB 7 | 0.00072 | 0.49757 | 6.85880 | 6.86295 | 3.24126 |
| TB 8 | 0.00004 | 0.50010 | -6.85958 | 6.86159 | 3.24540 |
| TB 9 | 0.00058 | 0.49913 | -6.86441 | 6.85852 | -3.21588 |

Additional Requested Measurements

Flatness of Endplates

TB side:	0.00066

Z+ side:	0.00025

Distance Endplate to Endplate:	8.23757

Parallelism of planes:	0.00499

Perpendicularity of Beamline Axis to

TB side:	0.00758

+Z side:	0.00086

Location of Z+ side End Plate Bore to TB side bore (TB SIDE BORE CSY)

X=	-0.00690
Y=	0.00315

# Text boxes

+Z

-Z

-Z

+Z

+X

+Y

+Y

+Z

+X

+Y