LCLS Tweeter Quadrupole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Carl Rago |
| Date: | Wednesday, August 09, 2006 |
| Work Order/Charge No.: | 92-4261-3 |
| Serial Number: | 001029 |
| URL of Fiducial Report: | \\Web002\www-group\met\Quality\FIDUCIAL REPORTS\LCLS TWEETER QUADS\TWEETER QUAD 4.pdf |

Part Set-up – Coordinate System Set-up

Planar Alignment

- Mid-Plane of the magnet

Spatial Alignment

- A plane on the top part of the magnet (red) by the Tooling Ball Sockets.

  - +X is from TB 3 & 4 side to TB 1 & 2 side.

“Z” Zero

- Mid-Plane of the magnet

“X” & “Y” Zero

- On both ends

  - Tangent point of each radii (4 on each end, 8 total).

    - Create a line between diagonal tangent points creates 2 lines.

      - Intersect the lines.

        - Creates a point on each end.

- Create a line of these 2 end points

  - This is the “X” & “Y” Zero, and Beamline or” Z” Axis.

| ISO | Front View |
| --- | --- |

Tooling Ball Measurements/Locations

Top of magnet; view from “+Y”

| Tooling Ball | FORM | DIAMETER | X | Y | Z |
| --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00069 | 0.49575 | 1.97152 | 6.74046 | 1.25205 |
| TB 2 | 0.00044 | 0.49663 | 1.96952 | 6.73871 | -1.26191 |
| TB 3 | 0.00058 | 0.49809 | -1.96839 | 6.73974 | -1.25148 |
| TB 4 | 0.00072 | 0.49587 | -1.96703 | 6.74232 | 1.25050 |

Pole Distances

| Pole | +Z side | -Z side |  |
| --- | --- | --- | --- |
| A-C | 1.69120 | 1.69084 | 0.00036 |
| B-D | 1.69164 | 1.19160 | 0.00004 |

# Text boxes

+Y

-Z

+Y

+X

+X

+Z

A

B

C

D