LCLS Undulator Quadrupole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Scott Anderson |
| Date: | Friday, May 30, 2008 |
| Work Order/Charge No.: | 88-1400-5 |
| Serial Number: | SLAC – 000706 / STAMPED – 29 |
| URL of Fiducial Report: | http://www-group.slac.stanford.edu/met/Quality/FIDUCIAL REPORTS/LCLS UNDULATOR QUADS/000706.pdf |

Part Set-up – Coordinate System Set-up

Spatial Alignment

- Geometric axis of the poles of the magnet.

Planar Alignment

- Plane on where Aluminum Base Rests

“Z” Zero

- Mid-plane of the magnet (middle of upstream and downstream ends).

“X” & “Y” Zero

- Geometric axis of the poles of the magnet.

| Front View | Side View |
| --- | --- |

Tooling Ball Measurements/Locations

Top View looking from +Y

Tooling Ball Adapter Cylinder Projected 1” Offset to the Tooling Ball Adapter Plane

| Tooling Ball | FORM | Cyl. Dia. | X | Y | Z | b |
| --- | --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00019 | 0.24999 | 3.37373 | 6.81522 | 1.49038 | 0.00010 |
| TB 2 | 0.00013 | 0.24998 | -3.37463 | 6.81633 | 1.50025 | 0.00011 |
| TB 3 | 0.00013 | 0.24989 | -3.38004 | 6.81239 | -1.50040 | 0.00008 |
| TB 4 | 0.00015 | 0.25001 | 3.36993 | 6.81162 | -1.50970 | 0.00009 |
| TB 5 | 0.00013 | 0.25000 | 6.58054 | 0.13204 | -0.00140 | 0.00013 |

Z-distance of Quad = 3.59526

Pole Tip Additional Data

*Data looking from Downstream End

| Side | Pole Dist A-C | Pole Dist B-D | Gap 1 | Gap 2 | Gap 3 | Gap 4 |
| --- | --- | --- | --- | --- | --- | --- |
| Downstream End (+Z) | 0.43533 | 0.43106 | 0.16211 | 0.15798 | 0.16151 | 0.15769 |
| Upstream End (-Z) | 0.43297 | 0.43395 | 0.15677 | 0.16162 | 0.16120 | 0.16008 |

| Side | A-Rad | B-Rad | C-Rad | D-Rad |
| --- | --- | --- | --- | --- |
| +Z | 0.21553 | 0.21554 | 0.21984 | 0.21555 |
| -Z | 0.21647 | 0.21756 | 0.21650 | 0.21639 |

| Side | J | K | L | M | N | O | P | Q |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| +Z | 0.08096 | 0.08091 | 0.07678 | 0.08180 | 0.08031 | 0.08070 | 0.07729 | 0.08056 |
| -Z | 0.08162 | 0.07800 | 0.08209 | 0.07721 | 0.07956 | 0.08008 | 0.08154 | 0.07959 |

Quadrupole Roll Angle

** Tolerance in milliradians is ± 5

Roll of the Base of Quad to Poles



milliradians

# Text boxes

+Y

+Z

+Y

+X

+Z

TB 5

+X