LCLS Undulator Quadrupole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Scott Anderson |
| Date: | Monday, April 28, 2008 |
| Work Order/Charge No.: | 88-1400-5 |
| Serial Number: | SLAC – 000708 / STAMPED – 09 |
| URL of Fiducial Report: | http://www-group.slac.stanford.edu/met/Quality/FIDUCIAL REPORTS/LCLS UNDULATOR QUADS/000708.pdf |

Part Set-up – Coordinate System Set-up

Spatial Alignment

- Geometric axis of the poles of the magnet.

Planar Alignment

- Plane on where Aluminum Base Rests

“Z” Zero

- Mid-plane of the magnet (middle of upstream and downstream ends).

“X” & “Y” Zero

- Geometric axis of the poles of the magnet.

| Front View | Side View |
| --- | --- |

Tooling Ball Measurements/Locations

Top View looking from +Y

Tooling Ball Adapter Cylinder Projected 1” Offset to the Tooling Ball Adapter Plane

| Tooling Ball | FORM | Cyl. Dia. | X | Y | Z | b |
| --- | --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00012 | 0.25005 | 3.36939 | 6.81429 | 1.50247 | 0.00012 |
| TB 2 | 0.00014 | 0.25000 | -3.38109 | 6.81090 | 1.50158 | 0.00023 |
| TB 3 | 0.00071 | 0.25003 | -3.37972 | 6.81366 | -1.50016 | 0.00039 |
| TB 4 | 0.00013 | 0.25007 | 3.37048 | 6.81475 | -1.49851 | 0.00017 |
| TB 5 | 0.00009 | 0.25021 | 6.57723 | 0.12245 | 0.00635 | 0.00011 |

Z-distance of Quad = 3.57993

Pole Tip Additional Data

*Data looking from Downstream End

| Side | Pole Dist A-C | Pole Dist B-D | Gap 1 | Gap 2 | Gap 3 | Gap 4 |
| --- | --- | --- | --- | --- | --- | --- |
| Downstream End (+Z) | 0.43265 | 0.43456 | 0.15838 | 0.15916 | 0.16142 | 0.16004 |
| Upstream End (-Z) | 0.43445 | 0.43363 | 0.15857 | 0.16124 | 0.15973 | 0.15856 |

Quadrupole Roll Angle

** Tolerance in milliradians is ± 5

Roll of the Base of Quad to Poles



milliradians

# Text boxes

+Y

+Z

+Y

+X

+Z

TB 5

+X