LCLS Undulator Quadrupole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Scott Anderson |
| Date: | Wednesday, September 10, 2008 |
| Work Order/Charge No.: | 88-1400-5 |
| Serial Number: | SLAC – 000709 / STAMPED – 30 (Re-Fiducialized) |
| URL of Fiducial Report: | http://www-group.slac.stanford.edu/met/Quality/FIDUCIAL REPORTS/LCLS UNDULATOR QUADS/000709_01.pdf |

Part Set-up – Coordinate System Set-up

Spatial Alignment

- Geometric axis of the poles of the magnet.

Planar Alignment

- Plane on where Aluminum Base Rests

“Z” Zero

- Mid-plane of the magnet (middle of upstream and downstream ends).

“X” & “Y” Zero

- Geometric axis of the poles of the magnet.

| Front View | Side View |
| --- | --- |

Tooling Ball Measurements/Locations

Top View looking from +Y

Tooling Ball Adapter Cylinder Projected 1” Offset to the Tooling Ball Adapter Plane

| Tooling Ball | FORM | Cyl. Dia. | X | Y | Z | b |
| --- | --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00015 | 0.25004 | 3.37569 | 6.81250 | 1.48547 | 0.00005 |
| TB 2 | 0.00015 | 0.24995 | -3.37627 | 6.81234 | 1.49546 | 0.00012 |
| TB 3 | 0.00020 | 0.24991 | -3.37990 | 6.81179 | -1.50388 | 0.00016 |
| TB 4 | 0.00016 | 0.24996 | 3.36877 | 6.81311 | -1.51257 | 0.00007 |
| TB 5 | 0.00013 | 0.25015 | 6.57795 | 0.13165 | 0.00589 | 0.00007 |

Z-distance of Quad = 3.59443

Pole Tip Additional Data

*Data looking from Downstream End

| Side | Pole Dist A-C | Pole Dist B-D | Gap 1 | Gap 2 | Gap 3 | Gap 4 |
| --- | --- | --- | --- | --- | --- | --- |
| Downstream End (+Z) | 0.43417 | 0.431103 | 0.16021 | 0.15799 | 0.16031 | 0.15797 |
| Upstream End (-Z) | 0.43209 | 0.43277 | 0.15477 | 0.16237 | 0.15702 | 0.16203 |

| Side | A-Rad | B-Rad | C-Rad | D-Rad |
| --- | --- | --- | --- | --- |
| +Z | 0.21859 | 0.21553 | 0.21558 | 0.21552 |
| -Z | 0.21607 | 0.21667 | 0.21602 | 0.21610 |

| Side | J | K | L | M | N | O | P | Q |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| +Z | 0.08152 | 0.08120 | 0.07676 | 0.08084 | 0.07938 | 0.07932 | 0.07867 | 0.07879 |
| -Z | 0.07839 | 0.08006 | 0.08197 | 0.07627 | 0.07850 | 0.08084 | 0.08153 | 0.07863 |

Quadrupole Roll Angle

** Tolerance in milliradians is ± 5

Roll of the Base of Quad to Poles



milliradians

# Text boxes

+Y

+Z

+Y

+X

+Z

TB 5

+X