LCLS Undulator Quadrupole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Scott Anderson |
| Date: | Wednesday, April 23, 2008 |
| Work Order/Charge No.: | 88-1400-5 |
| Serial Number: | SLAC – 000716 / STAMPED – 35 |
| URL of Fiducial Report: | http://www-group.slac.stanford.edu/met/Quality/FIDUCIAL REPORTS/LCLS UNDULATOR QUADS/000716.pdf |

Part Set-up – Coordinate System Set-up

Spatial Alignment

- Geometric axis of the poles of the magnet.

Planar Alignment

- Plane on where Aluminum Base Rests

“Z” Zero

- Mid-plane of the magnet (middle of upstream and downstream ends).

“X” & “Y” Zero

- Geometric axis of the poles of the magnet.

| Front View | Side View |
| --- | --- |

Tooling Ball Measurements/Locations

Top View looking from +Y

Tooling Ball Adapter Cylinder Projected 1” Offset to the Tooling Ball Adapter Plane

| Tooling Ball | FORM | Cyl. Dia. | X | Y | Z | b |
| --- | --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00028 | 0.25008 | 3.37219 | 6.81194 | 1.50229 | 0.00015 |
| TB 2 | 0.00009 | 0.25003 | -3.37856 | 6.81188 | 1.51024 | 0.00011 |
| TB 3 | 0.00010 | 0.24997 | -3.38458 | 6.81316 | -1.48614 | 0.00033 |
| TB 4 | 0.00013 | 0.24998 | 3.36924 | 6.81359 | -1.49577 | 0.00007 |
| TB 5 | 0.00021 | 0.25005 | 6.57954 | 0.14223 | 0.00966 | 0.00023 |

Z-distance of Quad = 3.59557

Pole Tip Additional Data

*Data looking from Downstream End

| Side | Pole Dist A-C | Pole Dist B-D | Gap 1 | Gap 2 | Gap 3 | Gap 4 |
| --- | --- | --- | --- | --- | --- | --- |
| Downstream End (+Z) | 0.43419 | 0.43206 | 0.16211 | 0.15871 | 0.15707 | 0.16070 |
| Upstream End (-Z) | 0.43181 | 0.43467 | 0.15760 | 0.16091 | 0.16086 | 0.16149 |

Quadrupole Roll Angle

** Tolerance in milliradians is ± 5

Roll of the Base of Quad to Poles



milliradians

# Text boxes

+Y

+Z

+Y

+X

+Z

TB 5

+X