LCLS Undulator Quadrupole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Scott Anderson |
| Date: | Tuesday, April 22, 2008 |
| Work Order/Charge No.: | 88-1400-5 |
| Serial Number: | SLAC – 000723 / STAMPED – 37 |
| URL of Fiducial Report: | http://www-group.slac.stanford.edu/met/Quality/FIDUCIAL REPORTS/LCLS UNDULATOR QUADS/000723.pdf |

Part Set-up – Coordinate System Set-up

Spatial Alignment

- Geometric axis of the poles of the magnet.

Planar Alignment

- Plane on where Aluminum Base Rests

“Z” Zero

- Mid-plane of the magnet (middle of upstream and downstream ends).

“X” & “Y” Zero

- Geometric axis of the poles of the magnet.

| Front View | Side View |
| --- | --- |

Tooling Ball Measurements/Locations

Top View looking from +Y

Tooling Ball Adapter Cylinder Projected 1” Offset to the Tooling Ball Adapter Plane

| Tooling Ball | FORM | Cyl. Dia. | X | Y | Z | b |
| --- | --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00009 | 0.24995 | 3.38125 | 6.81193 | 1.49169 | 0.00020 |
| TB 2 | 0.00012 | 0.24990 | -3.36694 | 6.81177 | 1.52114 | 0.00009 |
| TB 3 | 0.00015 | 0.25009 | -3.37843 | 6.81380 | -1.47707 | 0.00024 |
| TB 4 | 0.00013 | 0.25016 | 3.36771 | 6.81348 | -1.51067 | 0.00009 |
| TB 5 | 0.00015 | 0.25010 | 6.57815 | 0.13825 | -0.00718 | 0.00008 |

Z-distance of Quad = 3.58226

Pole Tip Additional Data

*Data looking from Downstream End

| Side | Pole Dist A-C | Pole Dist B-D | Gap 1 | Gap 2 | Gap 3 | Gap 4 |
| --- | --- | --- | --- | --- | --- | --- |
| Downstream End (+Z) | 0.43300 | 0.43226 | 0.15881 | 0.16078 | 0.15699 | 0.16119 |
| Upstream End (-Z) | 0.43240 | 0.43398 | 0.16130 | 0.15872 | 0.16207 | 0.15734 |

Quadrupole Roll Angle

** Tolerance in milliradians is ± 5

Roll of the Base of Quad to Poles



milliradians

# Text boxes

+Y

+Z

+Y

+X

+Z

TB 5

+X