LCLS Undulator Quadrupole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Scott Anderson |
| Date: | Thursday, January 25, 2008 (New TB Adapters) |
| Work Order/Charge No.: | 88-1400-5 |
| Serial Number: | SLAC – 000738 / STAMPED – 03 |
| URL of Fiducial Report: | http://www-group.slac.stanford.edu/met/Quality/FIDUCIAL REPORTS/LCLS UNDULATOR QUADS/000738.pdf |

Part Set-up – Coordinate System Set-up

Spatial Alignment

- Geometric axis of the poles of the magnet.

Planar Alignment

- Plane on where Aluminum Base Rests

“Z” Zero

- Mid-plane of the magnet (middle of upstream and downstream ends).

“X” & “Y” Zero

- Geometric axis of the poles of the magnet.

| Front View | Side View |
| --- | --- |

Tooling Ball Measurements/Locations

Top View looking from +Y

Tooling Ball Adapter Cylinder Projected 1” Offset to the Tooling Ball Adapter Plane

| Tooling Ball | FORM | Cyl. Dia. | X | Y | Z | b |
| --- | --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00015 | 0.24991 | 3.36640 | 6.81604 | 1.51747 | 0.00009 |
| TB 2 | 0.00012 | 0.24999 | -3.38665 | 6.81484 | 1.48719 | 0.00009 |
| TB 3 | 0.00020 | 0.24998 | -3.36973 | 6.80934 | -1.50853 | 0.00015 |
| TB 4 | 0.00015 | 0.24999 | 3.37889 | 6.811448 | -1.48112 | 0.00015 |
| TB 5 | 0.00015 | 0.24998 | 6.57942 | 0.14597 | 0.01553 | 0.00014 |

.

Pole Tip Additional Data

*Data looking from Downstream End

| Side | Pole Dist A-C | Pole Dist B-D | Gap 1 | Gap 2 | Gap 3 | Gap 4 |
| --- | --- | --- | --- | --- | --- | --- |
| Downstream End (+Z) | 0.43457 | 0.43394 | 0.16202 | 0.46073 | 0.16033 | 0.15980 |
| Upstream End (-Z) | 0.43697 | 0.43210 | 0.16082 | 0.15998 | 0.16089 | 0.15988 |

Quadrupole Roll Angle

** Tolerance in milliradians is ± 5

Roll of the Base of Quad to Poles



milliradians

Transformation

- Use TB locations 1, 2, & 3 to establish a Tooling Ball Plane this will be the Spatial Alignment.

- Use TB locations 1 & 2 to establish an axis where the line from TB 2 to TB 1 is the +X direction and will be the Planar Alignment.

- TB 1 will set zero in “X”, “Y”, & “Z”.

| Transformation | X | Y | Z |
| --- | --- | --- | --- |
| Displacement Vector | -85.70916 | -173.04175 | -38.47809 |
| Rotation Matrix | X | Y | Z |
| X | 0.99999 | 0.00018 | 0.00448 |
| Y | -0.00017 | 1.00000 | -0.00184 |
| Z | -0.00448 | 0.00184 | 0.99999 |

# Text boxes

+Y

+Z

+Y

+X

+Z

TB 5

+X