LCLS Undulator Quadrupole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Scott Anderson |
| Date: | Monday, June 16, 2008 |
| Work Order/Charge No.: | 88-1400-5 |
| Serial Number: | SLAC – 000744 / STAMPED – 14 |
| URL of Fiducial Report: | http://www-group.slac.stanford.edu/met/Quality/FIDUCIAL REPORTS/LCLS UNDULATOR QUADS/000744.pdf |

Part Set-up – Coordinate System Set-up

Spatial Alignment

- Geometric axis of the poles of the magnet.

Planar Alignment

- Plane on where Aluminum Base Rests

“Z” Zero

- Mid-plane of the magnet (middle of upstream and downstream ends).

“X” & “Y” Zero

- Geometric axis of the poles of the magnet.

| Front View | Side View |
| --- | --- |

Tooling Ball Measurements/Locations

Top View looking from +Y

Tooling Ball Adapter Cylinder Projected 1” Offset to the Tooling Ball Adapter Plane

| Tooling Ball | FORM | Cyl. Dia. | X | Y | Z | b |
| --- | --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00018 | 0.25007 | 3.37661 | 6.81346 | 1.49135 | 0.00010 |
| TB 2 | 0.00014 | 0.25002 | -3.37456 | 6.81502 | 1.51409 | 0.00013 |
| TB 3 | 0.00014 | 0.25003 | -3.38139 | 6.81463 | -1.48916 | 0.00007 |
| TB 4 | 0.00019 | 0.25010 | 3.36810 | 6.81202 | -1.50859 | 0.00023 |
| TB 5 | 0.00009 | 0.24999 | 6.57886 | 0.13666 | -0.00645 | 0.00021 |

Z-distance of Quad = 3.59438

Pole Tip Additional Data

*Data looking from Downstream End

| Side | Pole Dist A-C | Pole Dist B-D | Gap 1 | Gap 2 | Gap 3 | Gap 4 |
| --- | --- | --- | --- | --- | --- | --- |
| Downstream End (+Z) | 0.43246 | 0.43274 | 0.15836 | 0.16066 | 0.15858 | 0.16070 |
| Upstream End (-Z) | 0.43278 | 0.43353 | 0.16085 | 0.15830 | 0.16042 | 0.15908 |

| Side | A-Rad | B-Rad | C-Rad | D-Rad |
| --- | --- | --- | --- | --- |
| +Z | 0.21620 | 0.21658 | 0.21627 | 0.21616 |
| -Z | 0.21638 | 0.21722 | 0.21641 | 0.21632 |

| Side | J | K | L | M | N | O | P | Q |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| +Z | 0.07925 | 0.07953 | 0.08117 | 0.07884 | 0.07952 | 0.08011 | 0.08055 | 0.07933 |
| -Z | 0.07958 | 0.08044 | 0.08126 | 0.07977 | 0.08108 | 0.07952 | 0.07879 | 0.08084 |

Quadrupole Roll Angle

** Tolerance in milliradians is ± 5

Roll of the Base of Quad to Poles



milliradians

# Text boxes

+Y

+Z

+Y

+X

+Z

TB 5

+X