LCLS Undulator Quadrupole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Scott Anderson |
| Date: | Friday, June 13, 2008 |
| Work Order/Charge No.: | 88-1400-5 |
| Serial Number: | SLAC – 000745 / STAMPED – 26 |
| URL of Fiducial Report: | http://www-group.slac.stanford.edu/met/Quality/FIDUCIAL REPORTS/LCLS UNDULATOR QUADS/000745.pdf |

Part Set-up – Coordinate System Set-up

Spatial Alignment

- Geometric axis of the poles of the magnet.

Planar Alignment

- Plane on where Aluminum Base Rests

“Z” Zero

- Mid-plane of the magnet (middle of upstream and downstream ends).

“X” & “Y” Zero

- Geometric axis of the poles of the magnet.

| Front View | Side View |
| --- | --- |

Tooling Ball Measurements/Locations

Top View looking from +Y

Tooling Ball Adapter Cylinder Projected 1” Offset to the Tooling Ball Adapter Plane

| Tooling Ball | FORM | Cyl. Dia. | X | Y | Z | b |
| --- | --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00018 | 0.25008 | 3.38449 | 6.81203 | 1.48942 | 0.00012 |
| TB 2 | 0.00011 | 0.25009 | -3.36804 | 6.81314 | 1.52871 | 0.00009 |
| TB 3 | 0.00011 | 0.25016 | -3.37988 | 6.81555 | -1.47210 | 0.00006 |
| TB 4 | 0.00013 | 0.25007 | 3.37105 | 6.81355 | -1.50861 | 0.00009 |
| TB 5 | 0.00010 | 0.25001 | 6.58036 | 0.13356 | -0.00945 | 0.00013 |

Z-distance of Quad = 3.59461

Pole Tip Additional Data

*Data looking from Downstream End

| Side | Pole Dist A-C | Pole Dist B-D | Gap 1 | Gap 2 | Gap 3 | Gap 4 |
| --- | --- | --- | --- | --- | --- | --- |
| Downstream End (+Z) | 0.43233 | 0.43518 | 0.16015 | 0.16166 | 0.15994 | 0.16014 |
| Upstream End (-Z) | 0.43303 | 0.43313 | 0.15873 | 0.16232 | 0.16027 | 0.15898 |

| Side | A-Rad | B-Rad | C-Rad | D-Rad |
| --- | --- | --- | --- | --- |
| +Z | 0.21610 | 0.21912 | 0.21624 | 0.21606 |
| -Z | 0.21658 | 0.21646 | 0.21646 | 0.21668 |

| Side | J | K | L | M | N | O | P | Q |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| +Z | 0.08056 | 0.07923 | 0.08090 | 0.08048 | 0.07967 | 0.08001 | 0.08165 | 0.07938 |
| -Z | 0.08087 | 0.07934 | 0.07964 | 0.08021 | 0.07852 | 0.08161 | 0.08071 | 0.07940 |

Quadrupole Roll Angle

** Tolerance in milliradians is ± 5

Roll of the Base of Quad to Poles



milliradians

# Text boxes

+Y

+Z

+Y

+X

+Z

TB 5

+X