LCLS Undulator Quadrupole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Scott Anderson |
| Date: | Friday, June 13, 2008 |
| Work Order/Charge No.: | 88-1400-5 |
| Serial Number: | SLAC – 000749 / STAMPED – 07 |
| URL of Fiducial Report: | http://www-group.slac.stanford.edu/met/Quality/FIDUCIAL REPORTS/LCLS UNDULATOR QUADS/000749.pdf |

Part Set-up – Coordinate System Set-up

Spatial Alignment

- Geometric axis of the poles of the magnet.

Planar Alignment

- Plane on where Aluminum Base Rests

“Z” Zero

- Mid-plane of the magnet (middle of upstream and downstream ends).

“X” & “Y” Zero

- Geometric axis of the poles of the magnet.

| Front View | Side View |
| --- | --- |

Tooling Ball Measurements/Locations

Top View looking from +Y

Tooling Ball Adapter Cylinder Projected 1” Offset to the Tooling Ball Adapter Plane

| Tooling Ball | FORM | Cyl. Dia. | X | Y | Z | b |
| --- | --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00011 | 0.25007 | 3.37219 | 6.81218 | 1.50212 | 0.00010 |
| TB 2 | 0.00013 | 0.25004 | -3.37866 | 6.81411 | 1.49478 | 0.00042 |
| TB 3 | 0.00013 | 0.25015 | -3.37324 | 6.81744 | -1.49972 | 0.00011 |
| TB 4 | 0.00009 | 0.25009 | 3.37708 | 6.81298 | -1.49547 | 0.00009 |
| TB 5 | 0.00009 | 0.25003 | 6.57935 | 0.13276 | 0.00155 | 0.00016 |

Z-distance of Quad = 3.58839

Pole Tip Additional Data

*Data looking from Downstream End

| Side | Pole Dist A-C | Pole Dist B-D | Gap 1 | Gap 2 | Gap 3 | Gap 4 |
| --- | --- | --- | --- | --- | --- | --- |
| Downstream End (+Z) | 0.43473 | 0.43368 | 0.15972 | 0.16115 | 0.16130 | 0.15765 |
| Upstream End (-Z) | 0.43445 | 0.43211 | 0.15962 | 0.16038 | 0.16027 | 0.15908 |

| Side | A-Rad | B-Rad | C-Rad | D-Rad |
| --- | --- | --- | --- | --- |
| +Z | 0.21685 | 0.21685 | 0.21789 | 0.21684 |
| -Z | 0.21611 | 0.21608 | 0.21834 | 0.21604 |

| Side | J | K | L | M | N | O | P | Q |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| +Z | 0.08062 | 0.07831 | 0.07934 | 0.08045 | 0.07927 | 0.08201 | 0.07914 | 0.08069 |
| -Z | 0.08102 | 0.07906 | 0.08002 | 0.07966 | 0.07996 | 0.08099 | 0.07939 | 0.07925 |

Quadrupole Roll Angle

** Tolerance in milliradians is ± 5

Roll of the Base of Quad to Poles



milliradians

# Text boxes

+Y

+Z

+Y

+X

+Z

TB 5

+X