Orion Solenoid Magnet

Fiducialization

Inspector:   	Keith Caban

Customer: 		Dr. Dennis Palmer

Date: 			04/12/04

Charge #.:		12-1121-1

Coordinate System Set up:

- “Z” zero is set by taking the mid-plane of both end faces of the magnet.

- “X” and “Y zero is set by taking 2 inside diameters of the through cylinder.

  - This creates a single axis which is the simulated Beam Direction and is “X/Y” zero, with Z+ pointing in beam direction away (opposite direction) from the Return & Supply Fittings.

- Coordinate System is “clocked” by using a line on the lower negative side, this eliminates roll.

- All Coordinate Systems use the Right Hand Rule and Beam Direction is always Z+.

All Tooling Ball measurements are based on values within the above coordinate system, using ½” Diameter tooling balls with the 1” extension.

| Back View | Front View |
| --- | --- |
| 3-D | Side View |

Orion Solenoid Magnet

Set up Measurements:

| c –Z side | 0.00141 |
| --- | --- |
| c +Z side | 0.00083 |
| n –Z side | 2.89680 |
| Form –Z side | 0.00142 |
| n +Z side | 2.89949 |
| Form +Z side | 0.00135 |

Tooling Ball Measurements:

| Tooling Ball | FORM | DIAMETER | X | Y | Z |
| --- | --- | --- | --- | --- | --- |
| TB(1) | 0.00088 | 0.49728 | 9.49562 | 4.92627 | -4.12192 |
| TB(2) | 0.00016 | 0.49960 | 9.47168 | 4.94827 | 4.12003 |
| TB(3) | 0.00096 | 0.49798 | 3.22471 | 10.19835 | -4.11880 |
| TB(4) | 0.00025 | 0.49781 | 3.23857 | 10.19734 | 4.10670 |
| TB(5) | 0.00083 | 0.49761 | -3.22603 | 10.19936 | -4.11316 |
| TB(6) | 0.00075 | 0.49828 | -3.22988 | 10.20084 | 4.11064 |

# Text boxes

+Z

+Y

-Z

+Y

+X

+X

+Y

+Y

+X