SSRL SPEAR3 HORIZONTAL

SCRAPER FIDUCIALS

Inspector:  Keith Caban

Date:         September 17, 2004

Drawing #:SA-444-339-17

Charge #:  95-2702-0

Fiducial Coordinate System Set Up:

Horizontal Scraper:

- R 6.00’s cylindrical axis  is the Z direction

  - This sets the Z Axis as the Spatial Alignment

  - +Z points towards Tooling Ball “H(3)”

- The Center point of the R 6.00 cylinder.

  - This sets zero in Y , with +Y going towards R 6.00

- The O.D. of Flange sets Zero in X, Z.

- Axis is created by the Symmetry Points of 2 holes in line with the leak check grooves @ 90° from this Axis.

  - This is the planar alignment for X direction  where +X is Tooling Ball “V(3)”

*Used ½” Tooling Balls with 1” extensions.

NON FIDUCIAL DATA:

| FEATURE | ACTUAL | NOMINAL | ACT – NOM |
| --- | --- | --- | --- |
| RADIUS 6.00 | 5.83512 | 6.00 | -0.16488 |
| ANGLE of Cylindrical Axis to the X Axis | 89.77153 | 90 | -.22847 |

FIDUCIAL DATA:

| TB | FORM | DIA | X | Y | Z |
| --- | --- | --- | --- | --- | --- |
| H1 | 0.00041 | 0.49974 | 0.07647 | -160890 | -2.05492 |
| H2 | 0.00012 | 0.49930 | -1.98987 | -1.62657 | -0.04993 |
| H3 | 0.00065 | 0.49649 | 0.02733 | -1.60980 | 2.03408 |
| H4 | 0.00095 | 0.49640 | 2.10365 | -1.58525 | 0.01149 |

# Text boxes

+Y

+Z

+Z

+X