# MFD NON-COMPLIANCE REPORT # 4497

Date: 3/6/03

Job #: 17185

Sub ID:  2

Part Number: PF-342-120-97 Rev. 0

Part Description: Inner Pole Plate

Engineer: D. Kharakh

Ext: 3636

Reason For Non-compliance:

True Position of .003 MMC to –B- -C- is out of location by .018” sn002

See Profile plots attached

See inspection report

Engineer has decided to use part as is