XGamma Quad Magnets

Fiducialization

(GTF/LCLS)

Inspector:   	Keith Caban

Customer: 		Stephen Gierman

Date: 			02/19/04

Account No.:	9240025

Coordinate System Set up:

- “Z” zero is set by taking the mid-plane of the outer stainless steel surfaces.

- “X” and “Y zero is set by taking 2 diameters of a 1.500” Rod shimmed in inner housing, with the 2 circles, creating a line.

  - This creates a single axis which is the simulated Beam Direction and is “X/Y” zero.

- Coordinate System is “clocked” at a 45° angle in which the 4 threaded holes are directly opposite the Y corner axis.

- All Coordinate Systems use the Right Hand Rule and Beam Direction is always +Z.

All Tooling Ball measurements are based on values within the above coordinate system.

Units:		Inches

Features:	X, Y, Z, Diameter of Tooling Ball, & Form of Tooling Ball.

Magnets #2 & #4 Fiducials

| Fig 1.1 | Fig 1.2 |
| --- | --- |

MAGNET #2

| TOOLING BALL | X | Y | Z | DIAMETER | FORM |
| --- | --- | --- | --- | --- | --- |
| TB 1 | 7.68595 | 2.03484 | 0.01377 | 0.49363 | 0.00142 |
| TB 2 | 2.03127 | 7.68754 | 0.02945 | 0.49360 | 0.00130 |
| TB 3 | -2.03208 | 7.68698 | 0.02807 | 0.49323 | 0.00141 |
| TB 4 | -7.69045 | 2.03093 | 0.00620 | 0.49618 | 0.00106 |

MAGNET #4

| TOOLING BALL | X | Y | Z | DIAMETER | FORM |
| --- | --- | --- | --- | --- | --- |
| TB 1 | 7.68606 | 2.03343 | 0.03112 | 0.49802 | 0.00105 |
| TB 2 | 2.02950 | 7.69106 | 0.02812 | 0.49807 | 0.00103 |
| TB 3 | -2.03143 | 7.68549 | 0.01108 | 0.49671 | 0.00078 |
| TB 4 | -7.68845 | 2.02954 | -0.00682 | 0.49770 | 0.00079 |

Magnets #1 & #3 Fiducials

| Fig 1.3 | Fig 1.4 |
| --- | --- |

MAGNET #1

| TOOLING BALL | X | Y | Z | DIAMETER | FORM |
| --- | --- | --- | --- | --- | --- |
| TB 1 | -7.68427 | 2.02907 | 0.01888 | 0.49745 | 0.00055 |
| TB 2 | -2.02903 | 7.68904 | 0.00033 | 0.49888 | 0.00041 |
| TB 3 | 2.03052 | 7.68796 | -0.01825 | 0.49680 | 0.00088 |
| TB 4 | 7.69078 | 2.03071 | -0.02747 | 0.49497 | 0.00130 |

MAGNET #3

| TOOLING BALL | X | Y | Z | DIAMETER | FORM |
| --- | --- | --- | --- | --- | --- |
| TB 1 | -7.68939 | 2.02958 | -0.00891 | 0.49646 | 0.00078 |
| TB 2 | -2.02697 | 7.68591 | -0.00843 | 0.49583 | 0.00094 |
| TB 3 | 2.03519 | 7.68416 | -0.00847 | 0.49568 | 0.00095 |
| TB 4 | 7.68875 | 2.02904 | -.00762 | 0.49982 | 0.00015 |

Fig 1.5

| GAP MEAS. | Upstream X | Downstream X | Upstream Y | Downstream Y |
| --- | --- | --- | --- | --- |
| MAGNET 1 | 1.505 | 1.501 | 1.510 | 1.506 |
| MAGNET 2 | 1.509 | 1.505 | 1.505 | 1.510 |
| MAGNET 3 | 1.499 | 1.503 | 1.508 | 1.505 |
| MAGNET 4 | 1.505 | 1.509 | 1.506 | 1.509 |

# Text boxes

TB 4

TB 3

TB 2

TB 1

+Z

(Beam Dir)

UPSTREAM

+X

+Y

+X

+Y

Y GAP

UPSTREAM

DOWNSTREAM

X GAP

UPSTREAM

DOWNSTREAM

TB 1

TB 2

TB 4

TB 3

+Y

+Z

(Beam Dir)

UPSTREAM

+X

+X

+Y