MFD Vacuum Processing Request Form

Attach to MFD Job Input Form OR complete following fields:

Contact 	Ext. # Page # M/S

Date Submitted 		MFD File Number

Sketch or Part Number	Material(s): List all materials included in part(s)

Standard Bake Out Requested

- Temperature is 250°C

- Temperature ramp up and down is not controlled

- Bake Duration is 48hrs.

- Acceptance requires the following conditions must be met:

  - RGA Scan is taken at 150°C and the sum of all peaks above 44 AMU must be less than 1 X 10-11 Torr and no individual peak above 44 AMU may be greater than 5 X 10-12 Torr

  - The outgassing rate taken at room temperature must be less than 2 X 10-12 T-l/s/cm2

Modified Bake Out Requested

- Maximum Temperature °C

- Ramp rate up °C/hour

- Ramp rate down °C/hour

- Bake Duration

- Modified Acceptance Criteria

Additional Comments or Requests