[image: image2.png]ad_

S Y am \ Y

NATIONAL ACCELERATOR LABORATORY




[image: image2.png]                                                                                          [image: image3.png]FACET





SLAC Magnetic Measurement Plan and Traveler for the FACET Dipole Magnet 5D36 B203330T (style “5D36” with 2.5” gap)

Account Number to be charged: 9141202
This traveler covers the magnetic measurements plan for one of two water-cooled dipoles, of engineering type “5D36” with 2.5” gap, called 5D36, to be used as  B203330T in sector 20 in FACET. 

This is a very old dipole made in 1965 at an outside-SLAC magnet vendor. It is a long time since it was in a beamline and has been scanned for residual radioactivity – see the tag, is very little.
This magnet itself weighs about 5000 lb, and comes to you with 2 steel saw horses to sit on. In the beamline it will sit with its poletips vertical, you can orient it horizontally in your lab. It has yellow labels that indicate its beamline position in FACET   (B203330T). It has its own power supply. The beam enters the 5D36 dipole at the non terminal end. It should come to you with some steel magnetic mirror pieces over the terminal end ONLY.
1. Receiving Information:

	Received by (MM initials)
	

	Date received : (dd-mm-yyyy):
	         ____-____-2010

	Checked Magnet Number( 5D36),Optics name(B203330T)
	

	Magnet Engineer (Cherrill Spencer, ext 3474): verifies she has  inspected the magnet travelers and authorized the magnetic measurements 
	      Consider this signed by Cherrill Spencer on 28th January 2011

	If dipole does not have a barcode sticker then ask a Mech Tech to add one and write the 6 digit barcode number here: 
	


Preparation:

2. Power and LCW Connections: Unipolar PS  >=500A , >=40 V  required. Cosine-shaped ramp that mimics the ramps produced by an Ethernet-driven controller.

High pressure LCW system will be needed to cool this magnet in your measurement lab.

3. Magnet Orientation:

A beam direction arrow, should be visible on the top or side of the core, .Notice: This magnet has coil interconnects, power terminals and water fittings DOWNstream.  If there is no arrow, contact Cherrill Spencer, x3474. 

	Beam-direction arrow in place (initials):
	


4. Magnet Alignment:

This magnet has sockets for 4 tooling balls (TBs), they were fiducialized last week and their fiducialization data has been sent to Mike Gaydosh.
	Fiducialization of new TB complete         (Alignment Initials :)
	


BUT as you will not be using a stretched wire  apparatus to measure this magnet you do not need it to be aligned in your lab. 
5. Power leads and polarity:

                               [image: image1.png]ty

‘negative” horizontal
dipole

—

A~
- ? peam direstio” $ e_





The required polarity for 5D36 is to bend an electron beam DOWNWARDS.   Looking downstream the right hand pole should be a NORTH pole (as shown in the above figure if the dipole is sitting as a horizontally bending dipole.)
 Determine the electrical connection polarity (with supply outputting positive current) which produces the required magnetic field polarities. 

Place + and  - labels on or near the main power flags for the required polarity.

	Polarity established, power +/- labels applied
	


6. Cooling Water Flow Check:

This dipole has 8 water circuits, four per main coil and 2 temporary  manifolds that you will use. 

Adjust the differential pressure across the in and out manifolds (deltaP) to be about 150 psi.

Total water flow should be about 13.25 gpm. If you cannot adjust deltaP to 150 psi , first use 200 psi (or max), measure the flow, then make total flow 13.25 +/-0.2 gpm by adjusting the valve.

	        
	 Water flow [gpm]
	Check 

	Maximum  delta P _______   
	
	

	Adjusted to produce 13.25:


	13.25 or 
	    


7. Thermocouple placement and Thermal Test: 

This is an old magnet, been sitting almost outside for years and hasn’t seen current in a while, so the first time you power it (probably for this thermal test) please bring it up slowly (say, 50 amps increase at each step) and watch the temperature measurements for a minute as each current is reached. At 125 amps I predict an LCW temperature increase of 0.36 degrees C, if the increase is much higher than that then something is wrong and call Cherrill Spencer, ext 3474.
Place 5 thermocouples at these locations:

- any one of the coils – on the incoming LCW – on the outgoing LCW  -elsewhere on bench for ambient temperature measurement – on the steel core near any pole tip.

Arrange for temperatures to be read at least once every 3 minutes for the following thermal test:

After water flow has been checked on each circuit, run dipole to 280 A for three hours, reading the thermocouples, magnet current and voltage

	Confirm total LCW flow is 13.25 +/- 0.2 gpm
	Flow is ___________________gpm

	Name of thermal test file and date of test
	

	URL of webpage where datafile stored
	

	Final LCW temp (in) – Final LCW temp (out)
	=                       °C


If the rise in the LCW temperature is more than 4 degrees C at 280 amps, something is wrong; stop, call Cherrill Spencer at x3474 or her substitute. 

8. Initialization:

To be done only once, after successful completion of the water flow check and Thermal Test above. Use the cosine ramp with an equivalent linear rate of 15 A per second; minimum current of 5 A, maximum current of 400 A and pause time of 120 seconds at the minimum and maximum currents. Take the quad from 5 A to 400 A and back 7 times, ending up at 5 A

	Initialization completed: date , MMG initials:
	      ___/____/2011


9a. Standardisation parameters and detail: (For use in measurements below)

Ramp type: cosine shape

Minimum standardisation current: 10 amps

Maximum standardisation current : 280 amps

Pause time at minimum or maximum standardisation current: 30 seconds

Equivalent linear ramp rate for standardization process: 15 amps/second (increasing or decreasing)

Start at 10 A and ramp up to 280 A and down to 5 A THREE times, finishing at 10 amps.

Setting to desired current (BDES) : ramp UP to BDES at 5 amps/second

 All operating currents should be approached from below; except when I request a series of decreasing currents – these will be approached from a higher current.  When moving from one current to another for the strength or harmonics measurements use 5 amps/second and wait at least 30 seconds after the operating current has been reached before making any magnetic measurements. Put basic standardization details in all datafile comments.

	Confirm that magnet will ramp at 15 A/s (initials:)
	


9b. Multipole standardizations BEFORE any magnetic measurements

As this running at a lower current than it was designed for it needs some training along the specified standardization cycle before any magnetic measurements are taken. So take it through 10 whole standardizations (10x3=30 up and down cycles) before making any of the measurements listed below.
10. Poletip or gap Field Measurements with HALL probe, this is only type of magnetic measurement    I am asking for on 5D36.
The aperture is 2.5” tall. The pole is about 5” wide. 
Please use a Hall probe to measure the poletip field at 125.1 amps (after a regular standardization procedure). Please make a correction in the field value for the thickness of the Hall probe. I’m expecting 4.707 kG in the center of the gap.
	   Current                  Actual current
	 Poletip field :

	125.1 A nominal 
	

	   Date and inititals:
	 ____/____/2011


11. Magnetic Measurements of Field at center of gap and some other X positions with Hall Probe. 

Purpose of these measurements is to check the magnet is behaving properly. We will use measurements taken by you for FFTB use about 8 years ago to find the transfer functions (current required to reach a certain integrated strength).
Place your Hall probe as close to the X,Y, Z center of the gap as you can and measure the field at these currents:
25,  50, 75, 100, 125, 150, 175, 200, 225, 280 amps 

Then back to 25 amps from 280 amps, measure the B  at the same set of currents in reverse order. Remember to ramp at 5 amps/second and wait for 30 seconds after reaching the new current before starting a measurement. 

	Filename of. Strength Measurement
	                                       .ru___

	Date of measurement, initials 
	___/____/2011


Repeat this set of currents with the Hall probe at these X values (Y=0, Z~0):

(assuming is sitting with horizontal poles)

-5.0,-2.5, +2.5, +5.0 cm
	Filename of various X Measurements
	                                       .ru___

	Date of measurement, initials 
	___/____/2011


Achim Weidemann will look at this data just to see it makes some sense.  If you have time please repeat this set of currents at z=5cm and -5cm.
No need to run in reverse polarity.
	
	

	
	


12. Saving and Distributing the Magnetic measurement data

Please post all data as it is acquired at 

	http://www-group.slac.stanford.edu/met/MagMeas/MAGDATA/FACET_Sector_20/ etc 

____________________________________________________________________


in a subdirectory with the magnet name.

Please inform Magnet Engineer ( Cherrill Spencer at cherrill@slac.stanford.edu  (ext 3474). and /or achim@SLAC.Stanford.EDU,  wittmer@SLAC.Stanford.EDU of FACET) when the strength and harmonic measurement data are ready and again when this completed traveller is available for inspection and analysis, and for final approval .

 13.             .This section is to be completed by FACET Magnet Engineer.  

	Data been inspected & analyzed by Magnet Engineer 
	

	Based on Magnet Engineer analysis of data can confirm the pre-assigned beamline location (optics-deck name):
	   Optics beamline location:

	Dipole is released from Magnetics Group, signed and dated by FACET  Magnet Engineer:
	

	Remove the LCW from the coils and take the magnet to the designated storage area. Signed & dated by MMG personnel:
	


14.   This section is to be completed by FACET Beam Physicist (Walter Wittmer).  

	Checked that integrated strength data is satisfactory and have generated the polynomial function for the controlling database to set the magnet.

Nominal operating current is _______________amps.

Checked that this is within the capability of the assigned power supply _____________________
	      Signature and date:

	Checked that the multipole values at r=10mm are below the Physics Requirements tolerances (initial):
	 

	Magnet accepted for FACET (signed):
	

	Date accepted (Month-Day-Year):
	


When this traveler is completed, attach it to the fabrication traveler of this magnet and send these hardcopies to the designated person in Kathleen Ratcliffe’s group at MS 18 to be scanned into a pdf file. Then that electronic file will be stored in a place TBD and the hardcopies will be returned to sit with the magnet until it is installed in the beamline.

	Two travelers for this magnet been scanned into a pdf file; by_____________________
	Name of pdf file _________________________________

	Two travelers been returned to sit with magnet in its “waiting for installation” place.
	Signed & dated:


    Further Tasks to be done on this magnet will be recorded on its Fabrication Traveler.

  END OF FACET 5D36 (B203330T) MAGNET MAGNETIC MEASUREMENT PLAN AND TRAVELER   
FACET 5D36 (B203330T) Magnetic Measurement Plan and Traveller 1 of 6 


[image: image3.png]