[image: image2.png]ad_

S Y am \ Y

NATIONAL ACCELERATOR LABORATORY




[image: image2.png]                                                                                          [image: image3.png]FACET





SLAC Magnetic Measurement Plan and Traveler for the FACET Sector 10 Dipole Magnet B101EP, STYLE 1.0D71.4T SERIAL#1 REVISED 4Jan 12 
Account Number to be charged: 607301 SECOND REVISION 11Jan2012
This traveler covers the magnetic measurements plan for one of two water-cooled dipoles, of engineering type 1.0D71.4T, called B101EP, to be used as the first shared dipole in the side-by-side chicanes in sector 10 in FACET. Please call it B101EP in your MMG database.
This is a brand new dipole made in 2011, the core by an outside-SLAC magnet vendor and the coils at SLAC. Consequently it has no residual radioactivity.
This magnet itself weighs about 11340 lb, and comes to you without any base plate. The total length of the magnet is about 79 inches. It has yellow labels that indicate its lattice name in sector 10 of FACET   (B101EP). It will run in series with 5 other dipoles, 4 of which your group measured in 2002. They were called the LBCC chicane dipoles. How those magnets have been standardized and were measured dictates how the 2 new dipoles will be standardized. The beam enters the B101EP dipole at the terminal end, which is also the not-nose end. Note that the beam direction will be opposite to this in the second 1.0D71.4T dipole.
1. Receiving Information:

	Received by (MM initials)
	SDA

	Date received : (dd-mm-yyyy):
	         5-1-2012

	Checked Magnet Number( B101EP), serial #1 of 1.0D71.4T
	

	Magnet Engineer (Cherrill Spencer, ext 3474): verifies she has  inspected the magnet travelers and authorized the magnetic measurements 
	      

	If dipole does not have a barcode sticker then ask Magnet Engineer to add one and write the 6 digit barcode number here: 
	


Preparation:

2. Power and LCW Connections: Unipolar PS  >=925A , >35 V  required. Cosine-shaped ramp that mimics the ramps produced by an Ethernet-driven controller. All 6 dipoles in the two side-by-side chicanes will run in series off one power supply. The PS will be controlled by a BITBUS controller (this is a different style controller from the one used on the original sector 10 chicane).
This magnet uses the Multi-Contact style connectors which you tend not to use, so you will remove the four cylindrically shaped connectors and connect your power leads directly to the copper flags. In the beamline we will use two power input cables and two output power cables, you can use one each if you wish, they must be able to carry 925 amps each.
There are two trim coils, they are to be connected to a bipolar PS providing up to 15 amps.

The regular pressure LCW system will be needed to cool this magnet in your measurement lab.

 3. Magnet Orientation:

A beam direction arrow, should be visible on the top or side of the core, .Notice: This magnet has coil interconnects, power terminals and water fittings UPstream.  If there is no arrow, contact Cherrill Spencer, x3474. 

	Beam-direction arrow in place (initials):
	SDA


4. Magnet Alignment:

This magnet has sockets for 4 tooling balls (TBs), they will either have been fiducialized by the time the dipole reaches you or their fiducialization will be done in while the dipole is on your measuring table.
	Fiducialization of new TB complete         (Alignment Initials : JMC LG
	1/4/2012


Alignment crew should do the set up on measurement stand , and use the x,y,z axes of the whole magnet (as defined by TB measurements) such that the roll angle (angle of horizontal axis) is less than about 2.0  milliradian (or as small as possible), and pitch and yaw is also minimized.

Fiducialization data on location of tooling balls with respect to geometric axis when poles are aligned perfectly horizontally should be saved.

	Alignment completed, Roll Angle = 0  millirad (Alignment Initials :)
	LG. 1/4/2012


Your stretched wire system must be set up carefully too. Please make sure its wires are oriented so they are really horizontal relative to the horizontal axis defined by the alignment crew.  We will NOT use a rotating coil on these dipoles.

5. Power leads and polarity:

                               [image: image1.png]ty

‘negative” horizontal
dipole

—

A~
- ? peam direstio” $ e_





The required polarity for B101EP is negative horizontal dipole.   Looking downstream the upper- pole should be a NORTH pole as shown in the above figure. Don’t take the position of the terminals in the sketch literally, they will be upstream.    

 Determine the electrical connection polarity (with supply outputting positive current) which produces the required magnetic field polarities. 

Place + and  - labels on or near the main power flags for the required polarity.

	Polarity established, power +/- labels applied
	SDA


The trim coils should be hooked up to produce a north pole on the top pole.
6. Cooling Water Flow Check:

Each dipole has 8 water circuits, four per main coil and 2 built-in manifolds that you will use. 

Adjust the differential pressure across the in and out manifolds (deltaP) to be about 80 psi. Total water flow should be about 6.0 gpm. If you cannot adjust deltaP to precisely 80 psi , first get close to 80 psi, measure the flow, then make total flow 6.0 +/-0.2 gpm by adjusting the valve.

	        
	 Water flow [gpm]
	Check 

	Actual delta P _______   
	6.1 GPM
	SDA


7. Thermocouple placement and Thermal Test: 

This is a brand new magnet and has never seen any current, so the first time you power it (probably for this thermal test) please bring it up slowly (say, 50 amps increase at each step) and watch the temperature measurements for a minute as each current is reached. At 880 amps I predict an LCW temperature increase of 15.7 degrees C, if the increase is much higher than that then something is wrong and call Cherrill Spencer, ext 3474.
Place 5 thermocouples at these locations:

- any one of the coils – on the incoming LCW – on the outgoing LCW  -elsewhere on bench for ambient temperature measurement – on the steel core near any pole tip.

Arrange for temperatures to be read at least once every 3 minutes for the following thermal test:

After water flow has been checked on each circuit, run JUST the FIRST dipole to 925 A for three hours, reading the thermocouples, magnet current and voltage

	Confirm total LCW flow is 6.0 +/- 0.2 gpm
	Flow is 6.1 gpm

	Name of thermal test file and date of test
	Rtdat.ru1

	URL of webpage where datafile stored
	

	Final LCW temp (in) – Final LCW temp (out)
	= 13.9            °C


If the rise in the LCW temperature is more than 20 degrees C at 925 amps, something is wrong; stop, call Cherrill Spencer at x3474 or her substitute. 

8. Initialization:

To be done only once, after successful completion of the water flow check and Thermal Test above. Use the cosine ramp with an equivalent linear rate of 14 amps per second; minimum current of 20 A, maximum current of 925 A and pause time of 120 seconds at the minimum and maximum currents. Take the quad from 20 A to 925 A and back 10 times, ending up at 20 A.
	Initialization completed: date , MMG initials:
	      1/9/2012


9a. Standardisation parameters and detail: (For use in measurements below)

Ramp type: cosine shape

Minimum standardisation current: 20 amps

Maximum standardisation current : 925 amps

Pause time at minimum or maximum standardization current: 30 seconds

Equivalent linear ramp rate for standardization process: 14 amps/second (increasing or decreasing)
Start at 20 A and ramp up to 925 A and down to 5 A THREE times, finishing at 20 amps.

Setting to desired current (BDES) : ramp UP to BDES at 7 amps/second

 All operating currents should be approached from below; except when I request a series of decreasing currents – these will be approached from a higher current.  When moving from one current to another for the strength or harmonics measurements use 7 amps/second and wait at least 30 seconds after the operating current has been reached before making any magnetic measurements. Put basic standardization details in all datafile comments.

	Confirm that magnet will ramp at 14 A/s (initials:)
	SDA


9b. Multipole standardizations BEFORE any magnetic measurements

As this is a brand new magnet it needs some training along the specified standardization cycle before any magnetic measurements are taken. So take it through 10 whole standardizations (10x3=30 up and down cycles) before making any of the measurements listed below.
10. Poletip Field Measurements 

The aperture is 1.000” tall. The pole is 26cm wide not counting the chamfers. When it is convenient, maybe when the stretched wire is not in the aperture, please use a Hall probe to measure the poletip field  at 859 amps (after a regular standardization procedure), at the non-nose end. Please make a correction in the field value for the thickness of the Hall probe. 

	   Current                  Actual current
	 Poletip field : at x=0 and at x=10 cm 

	859 A nominal         858.96
	X = 0, B=  1.556,  x=10   B =  1.552  (Nose End)
X = 0, B= 1.565 T, x=10   B = 1.559    (non-Nose End)

	   Date and initials:
	1/18/2012


11. Magnetic Measurements of ∫Bdl and harmonics at various main currents. 

Purpose of these measurements is to find the transfer functions (current required to reach a certain integrated strength) especially for certain desired integrated strengths, and to check that the multipole harmonics do not exceed FACET requirements. Your stretched wire system must be set up carefully. Please make sure its wires are oriented so they are really horizontal relative to the horizontal axis defined by the alignment crew. For B101EP, the desired ∫Bdl is   -29.193 kGauss-m (predicting a nominal current of 859 A)
(a) Integrated Strength Measurements:

Measure the integrated strength ∫Bdl at these currents at the x=0 center of the poletips:
  50, 125, 200, 275, 350, 425, 500, 575, 650, 725, 800, 859, 875, 925 amps 

Then back to 50 amps from 925 amps, measure the ∫Bdl at the same set of currents in reverse order. Remember to ramp at 7 A/s and wait for 30 seconds after reaching the new current before starting a measurement. 

	Filename of Int. Strength Measurement
	Wiredat.ru2               

	Date of measurement, initials 
	1/9/2012 SDA


I will do a fit of ∫B.dl versus current for the increasing set of currents and see if my predicted current of 859 A for the dipole is close and calculate the actual operating current, which I will provide to you for the harmonics measurement.

(b) Measurement of effect of trim coils at various currents
Measure the integrated strength ∫Bdl at the main nominal current (or 859 amps)  at the x=y=0 center of the poletip with the trims at (i) -10, -8,-6,-4,-2 amps and then at (ii) 2,4,6,8,10 amps. Write down the signs of the main and trim currents.

	Filenames of Int. Strength Msts with trims on
	 Wiredat.ru4

	Date of measurements, initials 
	1/13/2012     SDA


(c) Measurement of integrated field at zero amps

Measure the integrated field at a zero-amp setting, reach it by following this procedure: standardize the magnet, finishing at 5 amps, then set it to nominal operating current, then ramp down to zero amps and simply switch off the power supply.  For this zero-amp in the main coil setting, please also find the trim-coil current necessary to null the integrated dipole field.

	Filename of Int. Strength Measurement at 0 A
	Wiredat.ru8

	Date of measurement, initials 
	1/18/2012     SDA


	Trim coil current necessary to null integrated field at zero main coil amps. Date, initials 
	Wiredat.ru8
1/18/2012     SDA


(d) Harmonics Measurements with trims off:
We need a precise measurement of the harmonics, which the customer has given the tolerances for at 10cm (ten cm); the quad, sextupole, octupole and decapole tolerances are all very tight and we will use the stretched wire setup to take the necessary data.  The poletip is 26 cm wide between the chamfers and I want integrated strength measurements at 41 different x positions (and y=0) and at the nominal operating current. From this data I will fit polynomials to the ∫B.dl versus x plots and I will calculate the multipoles. Run at 859 A if I haven’t given you a new current value.
Here are the 41 x positions in centimeters, x=0 is the width center of the poletip:
-12.0, -11.0, -10.0, -9.0, -8.0, -7.5,-7.0,-6.5, -6.0, -5.5,-5.0, -4.5, -4.0,-3.5, -3.0, -2.5,-2.0,-1.5, -1.0, -0.5, 0.0, 0.5, 1.0, 1.5, 2.0, 2.5, 3.0, 3.5, 4.0, 4.5, 5.0, 5.5, 6.0, 6.5, 7.0, 7.5, 8.0, 9.0, 10.0, 11.0, 12.0 cm
	Filename of Harmonics Measurement
	Wirevsx.ru3

	Date of measurement, initials 
	1/12/2012  SDA


(e) Harmonics Measurements with trims ON:
Repeat the measurements described in paragraph (d) with the trim coils on with 10 amps and with the polarity that adds to the main field.
	Filename of Harmonics Mst with TRIMS on
	Wirevsx.ru5

	Date of measurement, initials 
	1/13/2012  SDA


IF THE MULTIPOLE TOLERANCES ARE NOT MET I will have to machine the removable nosepiece to have some different shape and then the above measurements will be repeated. I may have some Hall Probe measurements done first.
12. Measurement of the field at various x and z at the nose-piece end with a Hall Probe

I will discuss the possibilities for this series of Hall Probe measurements with you in the new year. These Hall Probe measurements would be done after the stretched wire ones are completed.
13. Saving and Distributing the Magnetic measurement data

Please post all data as it is acquired at 

	http://www-group.slac.stanford.edu/met/MagMeas/MAGDATA/FACET_Sector_10/ B101EP


in a subdirectory with the magnet name.

Please inform Magnet Engineer ( Cherrill Spencer at cherrill@slac.stanford.edu  (ext 3474). and John Sheppard of FACET) when the strength and harmonic measurement data are ready and again when this completed traveller is available for inspection and analysis, and for final approval .

14.             .This section is to be completed by FACET Magnet Engineer.  

	Data been inspected & analyzed by Magnet Engineer 
	

	Based on Magnet Engineer analysis of data can confirm the pre-assigned beamline location (optics-deck name):
	   Optics beamline location:

	Dipole is released from Magnetics Group, signed and dated by FACET  Magnet Engineer:
	

	Remove the LCW from the coils and take the magnet to the designated storage area. Signed & dated by MMG personnel:
	


 15.   This section is to be completed by FACET Beam Physicist (John Sheppard).  

	Checked that integrated strength data is satisfactory and have generated the polynomial function for the controlling database to set the magnet.

Nominal operating current is _______________amps.

Checked that this is within the capability of the assigned power supply _____________________
	      Signature and date:

	Checked that the multipole values at r=10mm are below the Physics Requirements tolerances (initial):
	 

	Magnet accepted for FACET (signed):
	

	Date accepted (Month-Day-Year):
	


When this traveler is completed, attach it to the fabrication traveler of this magnet and send these hardcopies to the designated person in Kathleen Ratcliffe’s group at MS 18 to be scanned into a pdf file. Then that electronic file will be stored in a place TBD and the hardcopies will be returned to sit with the magnet until it is installed in the beamline.

	Two travelers for this magnet been scanned into a pdf file; by_____________________
	Name of pdf file _________________________________

	Two travelers been returned to sit with magnet in its “waiting for installation” place.
	Signed & dated:


    Further Tasks to be done on this magnet will be recorded on its Fabrication Traveler.

  END OF FACET B101EP (STYLE 1.0D71.4T SERIAL#1) MAGNET MAGNETIC MEASUREMENT PLAN AND TRAVELER   REVISION 1, dated 4th January 2012; REVISION 2, dated 11th January 2012
FACET B101EP (STYLE 1.0D71.4T SERIAL#1) Magnetic Measurement Plan and Traveller 6 of 7 


[image: image3.png]