## Test Procedure for the LBL BTA TQFB Q9 Quadrupole Magnets

(Revision 01)

This test procedure is intended to cover magnetic measurements for the LBL TQFB Q9 quadrupole magnets at SLAC.

Receiving:

The following information is to be noted upon beginning the measurements of the magnet by the SLAC MM group:

| Received by (initials): | SDA |
| --- | --- |
| Date put on Test stand (dd-mmm-yyyy): |  |
| Vendor Serial Number on the magnet: |  |
| Shipping damage (Y or N): |  |

Preparation:

Install magnet on test stand with rotating coil encoder on the upstream end of the magnet.

Procedures:

Ramp Rates and Wait Times:

Main Coil – 10 A/s, 20 seconds

Trim Coil – 2.5 A/s, 20 seconds

Linear Ramp

PROC01: Magnet Wash Cycle:

Ramp the magnet from 0 to 125 to 0 amps before starting measurements

Magnetic Measurements:

Enter URL of on-line magnetic measurements data:

http://www-group.slac.stanford.edu/met/MagMeas/MAGDATA/LBL/Quad/TQFB_0001_BTA_Q9

- Check that the quadrupole main and corrector connections are made so that a positive current produces fields as marked on the coils of the quadrupole for the quadrupole main , Bx trim and By trim.

- Confirm that the polarity for the main polarity markings on the magnet are correct:

| Quad Field polarity has been labeled correctly (initials): |  |
| --- | --- |

- Confirm that the polarity the Bx polarity markings on the magnet are correct:

| Bx trim lead connections are labeled correctly (initials): |  |
| --- | --- |

- Confirm that the polarity the By polarity markings on the magnet are correct:

| By lead connections are labeled correctly (initials): |  |
| --- | --- |

- Measure the inductance and resistance of the Quadruple coils:

| Inductance of Main coils at 100 Hz (mH): | 2.4417 mH |
| --- | --- |
| Resistance of Main coils (): | 0.1900 |
| Main Coil Temperature | 20.4 °C |

- Measure the inductance and resistance of the Bx trim coils:

| Inductance of Bx trim coils at 100 Hz (mH): | mH |
| --- | --- |
| Resistance of Bx trim coils (): |  |

- Measure the inductance and resistance of the By trim coils:

| Inductance of By trim coils at 100 Hz (mH): | mH |
| --- | --- |
| Resistance of By trim coils (): |  |

- Determine the number of current cycles at each current setpoint for main quadrupole harmonic integral reproducibility to < 3 units. Follow PROC01, then measure the length-integrated field gradient Gdl, at main current = 45 A for as many cycles as needed until main quadrupole harmonic integral reproducibility to < 3 units. Note the number of cycles needed.  This number of cycles will be used for each current setpoint.

| Rotating coil name | 1DQB26 |
| --- | --- |
| Rotating coil radius (m): | 0.012431 |
| Harmonics data file name: | Hardat.ru, Harplt.ru |
| Filename of Gdl data: | Strdat.ru, strplt.ru |
| # of Current Cycles need for < 3 unit reproducibility |  |

- Perform the wash cycle (PROC01), then measure the length-integrated field gradient and the field harmonics, for quadrupole current settings 55, 90, 125, 150, 170 and 180 amps using the number of cycles in set 8.

| Harmonics data file name: | Hardat.ru, Harplt.ru |
| --- | --- |
| Filename of Gdl data: | Strdat.ru, strplt.ru |

- Set the main to the nominal calibrated current of 122.79 A, (∫Gdl  = -13.71 T).  Measure the harmonics with the Bx corrector trim at 0, +10, +25, -10, -25 and 0 amps.

| Harmonics data file name: | Hardat.ru, Harplt.ru |
| --- | --- |
| Filename of Gdl data: | Strdat.ru, strplt.ru |

- Set the main to the maximum calibrated current of 94.28 A, (∫Gdl  = -10.551 T).  Measure the harmonics with the Bx corrector trim at 0, +10, +25, -10, -25 and 0 amps.

| Harmonics data file name: | Hardat.ru, Harplt.ru |
| --- | --- |
| Filename of Gdl data: | Strdat.ru, strplt.ru |

- Set the main to the nominal calibrated current of 122.79 A, (∫Gdl  = -13.71 T).  Measure the harmonics with the By corrector trim at 0, +10, +25, -10, -25 and 0 amps.

| Harmonics data file name: | Hardat.ru, Harplt.ru |
| --- | --- |
| Filename of Gdl data: | Strdat.ru, strplt.ru |

- Set the main to the maximum calibrated current of 94.28 A, (∫Gdl  = -10.551 T).  Measure the harmonics with the By corrector trim at 0, +10, +25, -10, -25 and 0 amps.

| Harmonics data file name: | Hardat.ru, Harplt.ru |
| --- | --- |
| Filename of Gdl data: | Strdat.ru, strplt.ru |

- Upon completion of tests, send link to all data to Martin Johansson .  Acceptance of magnet should be given by email to Magnetic Measurements Group.

| Acceptance given by Martin Johansson, date |  |
| --- | --- |