
[image: image1]

SLAC Traveler for HE-SXR Undulator #29A804-003 HE 
(Dover bench)
Dataset 001
This traveler is intended to cover preparation and magnetic measurements of the SXR undulator segments in the Dover bench.   

Preparation:
· Remove all protective covers 
· Clean up the granite

· Check linear potentiometers. Install brackets if necessary
· Label modules from 1 to 3
· Label undulator (HE-SXU-0XX)

	Vendor serial number from undulator label
	

	Engineer (initials) / date(mm-dd-yyyy):
	


· Call controls group to connect and check the control system
· Call controls group and/or MFD to set hard stops
· Call alignment crew to align the undulator to the bench
· Cycle gap, ~5000 gap changes (between 7.2mm and 33 mm) if it has not been done before
 Mechanical measurements:

· Open gap to 10mm (tuning gap). Check gap straightness with the touch probe every 10 poles, correct if necessary by replacing magnet keepers’ shims
· Correct encoders’ offsets to make jaws parallel and level

· Set half-gap encoders’ offsets at 10mm gap to read 5.0000mm

· Check compensation spring stack gaps (refer to nominal values)

· Measure every pole at 10mm gap, correct pole positions to ±20µm, if necessary

· At 10mm gap set the midplane encoders’ offsets, so the encoders read 5.0000mm

· Remove touch probe. Install and align the Hall probe 

· Find the third pole z-position                                  Z3 =
	Engineer (initials) / date (mm-dd-yyyy):
	


Pre-tuning
Follow the check list below, consult with the tuning test plan (LCLS-TN-22-09). 
Pre-tuning process checklist:

	Engineer (initials):
	

	Pre-tuning started, date (mm-dd-yyyy):
	

	· Undulator magnetically aligned at tuning Gap = 10mm (adjust encoder offsets if necessary), results recorded into the table below

	· MA shift vs Gap measured with Hall probe for gaps = 7.2,8.5,10,13,17,25, and 33 mm

	· Fields measured for central tuning x = ± 2mm at tuning gap=10mm

	· Downstream PM measured

	· Dover bench OFF, Long coil set, HP3458 and ESP302 ON

	· Optical measurement of the coil

	· Field integrals measured for tuning trajectories 


 

	· Coil and stages in park position 

· 

	· Shim table prepared. Gap opened to 180mm

	· Undulator trajectory shimming completed

	· Dover bench ON and re-initialized

	· Trajectories and phase errors checked. (If additional shimming is needed, repeat Hall probe measurements, coil measurements and tuning)

	· Trajectories and phase errors tuned to tolerances

	· Fields measured for tuning ends

	· Dover bench OFF, Long coil set, HP3458 and ESP302 ON

	· Field integrals measured for tuning first and second field integrals

	· Shim table prepared. Gap opened to 180mm. Coil and stages in park position



	· Field integrals tuning done

	· Long coil set back, HP3458 and ESP302 ON

	· Undulator field integrals checked. If additional shimming is needed, repeat coil measurements and tuning

	· Field integrals shimming completed

	· Coil and stages set in park position

	· Dover bench ON and re-initialized

	· MA position vs Y checked

	· Trajectories and phase errors checked and within the tolerances 

	· Yaw and Pitch checked

	· Controller Configuration saved


	Engineer (initials) / date (mm-dd-yyyy):
	


The following information is to be noted by an engineer:

	Average X (m):
	

	Average Y (m):
	

	Final Pitch (µrad):
	

	Final Yaw (µrad):
	

	Magnetic axis position X (m):
	

	Magnetic axis position Y (m):
	

	Coil X-position US offset (m)
	

	Coil Y-position US offset (m)
	

	Coil X-position DS offset (m)
	

	Coil Y-position DS offset (m)
	

	Downstream RM X offset (m):
	

	Downstream RM Y offset (m):
	

	Downstream RM Z offset (m):
	


Prepare the undulator to be moved:
· Open gap to 180 mm, hit emergency stops, power OFF the controller. Disconnect the network and power cable, and coil them up
· Remove linear potentiometer brackets, if installed
· Install all protective covers

· The undulator is ready to be moved
	Engineer (initials) / date (mm-dd-yyyy):
	


URL for data:

	V:\MET\MagServe\MagData\ DOVER BENCH\LCLS-II-HE\Undulator\HE-SXU_003\DATASET0001



Comments:
	


	Undulator pre-tuning completed (signed):
	


[image: image1][image: image2.png]


