## SLAC Traveler for HE-SXR Undulator #29A804-020 HE

## in MMF

Dataset 001

This traveler is intended to cover preparation, magnetic measurements, and fiducialization of the SXR undulator segments at MMF.

Preparation:

- Remove all protective covers. Install wooden platform between the undulator and granite.

- Remove plastic cover. Clean up the granite.

- Attach temperature sensors.

- Attach ground wire for capacitive sensors.

- Check linear potentiometers. Install brackets if necessary.

- Label modules from 1 to 3.

- Install coil supports.

| Vendor serial number from undulator label: |  |
| --- | --- |
| Technician (initials): |  |
| Date/Time: |  |

- Call controls group to connect and check the control system. Check completed.

- Call alignment crew to align the undulator to the bench. Alignment done.

- Cycle gap, ~5000 gap changes (between 7.2mm and 33 mm)

Mechanical measurements:

- Open gap to 10mm (tuning gap). Check gap straightness with the touch probe, correct if necessary by replacing magnet keepers’ shims.

- Correct encoders’ offsets to make jaws parallel and level.

- Set half-gap encoders’ offsets at 10mm gap to read 5.0000mm.

- Check compensation spring stack gaps (refer to nominal values).

- Measure every pole at 10mm gap, correct pole positions to ±20µm, if necessary.

- At 10mm gap set the midplane encoders’ offsets, so the encoders read 5.0000mm.

- Remove touch probe. Install the Hall probe, align.

- Find the third pole z-position.                                  Z3 =

| Engineer(initials)/date and time: |  |
| --- | --- |

Tuning.

Follow the check list below, consult with the tuning test plan (LCLS-TN-22-09).

Tuning process checklist:

| Engineer (initials): |  |
| --- | --- |
| Tuning started, Date (mm-dd-yyyy): |  |
| - Undulator magnetically aligned at tuning Gap = 10mm (adjust encoder offsets if necessary), results recorded into the table below. |  |
| - MA shift vs Gap measured with Hall probe for gaps = 7.2,8.5,10,13,17,25, and 33 mm |  |
| - Fields measured for central tuning x = ± 2mm at tuning gap=10mm |  |
| - Upstream PM measured |  |
| - Kugler bench OFF, Long coil set, PDI and XPS ON |  |
| - Field integrals measured for tuning trajectories |  |
| - Coil and stages in park position, PDI and XPS OFF, the coil removed for shimming. |  |
| - Shim table prepared. Gap opened to 180mm. |  |
| - Undulator trajectory shimming completed |  |
| - Kugler bench ON and re-initialized |  |
| - Trajectories and phase errors checked. (If additional shimming needed, repeat Hall probe measurements, coil measurements and tuning). |  |
| - Trajectories and phase errors tuned to tolerances. |  |
| - Fields measured for tuning ends. |  |
| - Kugler bench OFF. Long coil set, PDI and XPS ON |  |
| - Field integrals measured for tuning first and second field integrals. |  |
| - Shim table prepared. Gap opened to 180mm. The coil removed for shimming. |  |
| - Field integrals tuning done. |  |
| - Long coil set back, PDI and XPS ON. |  |
| - Undulator field integrals checked. If additional shimming needed, repeat coil measurements and tuning. |  |
| - Field integrals shimming completed. |  |
| - Field integrals measured for finding commissioning gap. |  |
| - Coil and stages set in park position, PDI and XPS OFF, the coil removed. |  |
| - Kugler bench ON and re-initialized. |  |
| - MA position vs Y checked. |  |
| - Trajectories and phase errors checked and within the tolerances. |  |
| - Hall probe measurements to find commissioning gap done. (K=6.08) |  |
| - Phase matching done. |  |
| - Coil installed; optical measurements of coil position done |  |
| - Field integrals vs gap checked. |  |
| - Parameters file checked before final |  |
| - Final scan done, coil |  |
| - Coil and stages set in park position. PDI and XPS OFF, the coil removed. |  |
| - Gap clearance checked. Gap =7.350mm |  |
| - Kugler bench ON and re-initialized. |  |
| - Yaw and Pitch checked |  |
| - Reference magnet measured. |  |
| - MA position vs Y checked. Gap = 10mm |  |
| - Half-gap encoders checked; Gap=10mm |  |
| - Commissioning gap checked. |  |
| - Parameters file checked before final |  |
| - Final scan done, Hall probe. |  |
| - PMs measured for fiducialization, file with results created. |  |
| - Fiducialization done. |  |
| - Touch probe measurements at 10mm gap done. |  |
| - Touch probe measurements with taper +100µm, +300µm, -100µm done. |  |
| - Kugler bench OFF |  |
| - Capacitive sensor measurements, poles #10, 15, 101, 106 done. |  |
| - Tuning Evaluation traveler created. |  |
| - Controller Configuration saved. |  |

The following information is to be noted by an engineer:

| Average X (m): |  |
| --- | --- |
| Average Y (m): |  |
| Final Pitch (µrad): |  |
| Final Yaw (µrad): |  |
| Magnetic axis position X (m): |  |
| Magnetic axis position Y (m): |  |
| Coil X-position (m) |  |
| Coil Y-position (m) |  |
| Commissioning Gap (mm) |  |
| Reference magnet measurement (T) |  |
| US Half-Gap Encoder |  |
| DS Half-Gap Encoder |  |

Fiducialization:

The following information is to be noted by an engineer:

| Engineer (Initials): |  |
| --- | --- |
| Upstream RM X offset (m): |  |
| Upstream RM Y offset (m): |  |
| Downstream RM X offset (m): |  |
| Downstream RM Y offset (m): |  |

Have an alignment crew to measure offsets between RM and the undulator fiducials.

| Surveyor (initials): |  |
| --- | --- |
| Date (mm/dd/yyyy): |  |

Prepare the undulator for the move into the storage area:

- Power OFF the controller. Disconnect the network cable.

- Remove linear potentiometer brackets, if installed.

- Remove long coil brackets.

- Remove temperature sensors.

- Remove ground wire for capacitive sensors.

- Install all protective covers. Remove wooden platform between the undulator and granite.

- The undulator is ready for move into the storage area.

| Technician (initials) / date: |  |
| --- | --- |

URL for data:

V:\MET\MagServe\MagData\LCLS-II-HE\Undulator\HE-SXU_020\DATASET0001\Final Results\

| Undulator tuning completed (signed): |  |
| --- | --- |