SXR Reference re-measurement results:  
8/3/2023
	DS
	Date
	Temp. (ºC)
	I1Y (µTm)
	I2Y (µTm2)
	I1X (µTm)
	I2X (µTm2)
	Phase RMS (º)
	K           (at 20ºC)
	(K/K
(vrt. DS01)
	Ref.
( (B/B)
	HP calibration

	0001
	09/17
	20.0
	-22
	-50
	+1
	+37
	3.8
	3.347130
	-
	-
	12/1/17

	0002
	1/18
	20.01
	-23
	-55
	-15
	-24
	3.8
	3.347029
	-3 · 10-5
	-4 · 10-5
	1/31/18

	0003
	5/18
	19.93
	-18
	-54
	-15
	-20
	3.8
	3.347334
	+6 · 10-5
	+8 · 10-5
	5/11/18

	0004
	7/18
	20.00
	-20
	-53
	-10
	-20
	3.8
	3.347382
	+8 · 10-5
	+8 · 10-5
	7/11/18

	0005
	9/18
	19.99
	-23
	-54
	-7
	-25
	3.8
	3.347252
	+4 · 10-5
	+8 · 10-5
	9/25/18

	0006
	12/18
	19.96
	-19 (-18)
	-36
	-18 (-23)
	-27
	3.7
	3.347402
	+8 · 10-5
	+7 · 10-5
	12/12/18

	0007
	4/19
	19.92
	-20
	-49
	-21
	-55
	3.8
	3.347493
	+1.0· 10-4
	+8 · 10-5
	4/3/19

	0008
	1/20
	19.88
	-23
	-54
	-39
	-89
	3.8
	3.347544
	+1.2· 10-4
	+1.3· 10-4
	1/16/20

	008
	8/20
	19.78
	-23
	-54
	-39
	-89
	3.8
	3.347249
	+4 · 10-5 
	+9 · 10-5
	8/17/20

	0009
	3/21
	19.87
	-23
	-54
	-39
	-89
	3.9
	3.347270
	+4 · 10-5
	+4 · 10-5
	3/10/21

	0010
	7/21
	19.92
	-23
	-54
	-39
	-89
	3.9
	3.347475
	+1.0· 10-4
	+4.7· 10-5
	3/10/21

	DS
	Date
	Temp. (ºC)
	I1Y (µTm)
	I2Y (µTm2)
	I1X (µTm)
	I2X (µTm2)
	Phase RMS (º)
	K  (at 20ºC)
	(K/K
(vrt. DS01)
	Ref.

( (B/B)
	HP calibration

	0011
	10/22
	19.88
	-23
	-54
	-39
	-89
	3.8
	3.347184
	+1.6· 10-5
	+3.4· 10-5
	10/24/22

	0012
	1/23
	19.90
	-23
	-54
	-39
	-89
	3.8
	3.347168
	+1.1· 10-5
	+0.8· 10-5
	1/11/23

	0013*
	8/23
	19.93
	-19(-17)
	-29
	-17(-24)
	-5
	3.7
	3.346941
	-5.6· 10-5
	+4· 10-5
	7/20/23

	HE-SXUs

	DS
	Date
	Temp. (ºC)
	I1Y (µTm)
	I2Y (µTm2)
	I1X (µTm)
	I2X (µTm2)
	Phase RMS (º)
	K  (at 20ºC)
	(K/K
(vrt. DS01)
	Ref.

( (B/B)
	HP calibration

	0014
	10/23
	19.96
	-19(-17)
	-29
	-17(-24)
	-5
	3.7
	3.347048
	-2.4· 10-5
	+3.4· 10-5
	7/20/23

	0015
	5/24
	19.95
	-19
	+4
	-18
	-8
	3.8
	3.346849
	-8.4· 10-5
	+0.8· 10-5
	5/7/24

	0016
	12/24
	19.83
	-33
	-55
	-20
	-12
	3.7
	3.347203
	2.2· 10-5
	+1.8· 10-5
	12/3/24

	0017
	4/25
	19.88
	-24
	-7
	-11
	+20
	3.7
	3.347167
	1.1· 10-5
	+1.1· 10-4
	12/3/24

	0018
	7/25
	19.82
	-25
	-31
	-8
	+26
	3.7
	3.347327
	5.9· 10-5
	+1.0 · 10-4
	7/23/25

	0019
	11/25
	19.86
	-31
	-10
	-12
	+26
	3.8
	3.34694
	-5.7· 10-5
	+8 · 10-5
	11/13/25


Avrg. K =  3.347270 ±5∙10-5
(-18) – Vibrating Wire measurement; * - coil position adjusted in Z to make it centered on undulator.
[image: image1.png]SXR Reference K change in time

—@— Re-measurement % Probe re-calibration —@— Corrected for Probe change

-2
4/25/2017 2/19/201812/16/20180/12/2019 8/7/2020 6/3/2021 3/30/2022 1/24/202311/20/20239/15/2024 7/12/2025
Date




[image: image2.png]Rel. change

Change w.r.t. first Dataset

1.5E-04

1.0E-04

5.0E-05

0.0E+00

-5.0E-05

-1.0E-04

—@—Hall probe - NMR rel. difference

DS#

—8—Und. Krel. difference

20




HE started








