[image: image3.jpg][ (¥

Lo\ >






Magnetic Measurements for the LCLS

XLEAP Chicane Dipole Magnets

Contact: D. Bohler  (March 2018)


Account number: 17838 –A8001
This traveler is intended to prescribe the magnetic measurements plan of the four XLEAP chicane dipole magnets (1.575D14-C). These rectangular C-type dipole magnets are 14 inches long with 8-mm gap height and have MAD names of: “BCXXL1”, “BCXXL4”, “BCXXL4”, and “BCXXL4”, with each magnet including both main coils and ~1% trim coils.  The main and trim coils are each designed for excitation at up to 13 Amperes and 6 Amperes respectively, requiring no water cooling and no over-temperature trip relays.

Receiving:

The following information is to be noted upon receipt of the magnets:

	Received by (initials):
	SDA

	Date received (dd-mmm-yyyy):
	5/31/2018

	Magnet serial number, if available:
	SLM3_04

	SLAC barcode number:
	2045533

	Labeled w/ MAD Name:
	BCXXL4


Preparation:
A beam direction arrow, with text “Beam Direction”, is to be applied to the connector side of the magnet on the iron between terminal strips.  With the C-magnet’s open-gap side facing away from the observer (terminal strips toward observer), the beam direction is to the left.  If this is not done please contact D. Bohler (3295).

	Beam-direction arrow properly in place (initials):
	 SDA


Initial Setup:

1) Verify that the magnets are complete and undamaged, including wiring connections.

	Incoming inspection OK (initials):
	SDA


Enter URL of on-line magnetic measurements data (please modify, append, or correct this URL as necessary):

	http://www-group.slac.stanford.edu/met/MagMeas/MAGDATA/LCLS/dipole/BCXXL4



2) Mark each magnet as BCXXL1, BCXXL4, BCXXL3, or BCXXL4.  By choosing the magnet names (locations) initially, they will be tested in their proper polarities, since two are to be positive and two negative.

	Magnet marked as (BCXXL1, BCXXL2, BCXXL3, or BCXXL4):
	BCXXL4


3) Determine the main-coil connection polarity (with MCOR30 main supply outputting positive current) which produces a “positive” (P) field polarity for BCXXL1 and BCXXL4 (below top), but a “negative” (N) field polarity for BCXXL2 and BCXXL4 (below left), as shown below:

[image: image1.png]+V  “pecifi
Y “positive” horizontal
dipole

ction

o=
g > pearn diré



[image: image2.png]ty

‘negative” horizontal
dipole

—

A~
- ? peam direstio” $ e_





Figure 1.  Dipole magnet sketch indicating field polarities.  Note that BCXXL1 and BCXXL4 are “positive” (top), while BCXXL2 and BCXXL4 are “negative” (bottom).  The C-magnet’s open-gap side (not shown here) is the side seen facing the observer.

4) Mark the polarity near the main magnet leads with clear “+” and “” labels as shown above.

	Main coil polarity chosen from Fig. 1 is (P or N):
	P


5) Also mark the trim leads with clear “+” and “” labels such that, with the trim supply outputting positive current, the trim coil increases the absolute value of the magnetic field established by the main coil.  This will set the trim polarity as “positive” (P) for BCXXL1 and BCXXL4 and “negative” (N) for BCXXL2 and BCXXL4, as described in LCLS-I PRD 1.1-010.

	Trim coil polarity chosen from Fig. 1 is (P or N):
	P


6) Connect the main magnet terminals (not the trims) in the correct polarity as established above, to a bipolar MCOR30 power supply with maximum current Imain = ±13 A.  Leave the trim coil disconnected for now.

7) In an environment with ambient temperature of about 20(C (68(F), excite the magnet’s main coils at 13 A for ~5 hours to warm it up (verify this is steady-state temp. and record value).

	Ambient temperature (°C):
	22.4

	Final magnet top steel surface temperature (°C):
	31.9

	Final magnet top coil surface temperature (°C):
	40.2


Magnetic Measurements
8) Standardize magnet using a three-linear cycle, starting from zero to +13 A, then to 13 A, and finally back to zero, through three of these full cycles, and ending again at zero, all with a flat-top pause time (at each setting of 13 and +13 A) of 5 seconds.  Use a ramp rate of 2.0 A/sec, if possible, but record the ramp rate used.

	Bipolar standardization complete (initials):
	SDA

	Ramp rate used (A/sec):
	2


9) Maintaining this three-linear cycle history, and with the trim coils still not powered, measure the length-integrated vertical dipole field, (Bydl, from -13 to +13 A in 0.5-A steps (53 ‘up’ measurements with at least a 2-sec pause at each setting).  Then, still maintaining the cycle history, measure (Bydl back down from +13 A to 13A in 0.5-A steps (53 ‘down’ measurements).
	Filename & run number of  (Bydl up, down, up data:
	Wiredat.ru1, wireplt.ru1


10) Repeat the bipolar standardization of step #8 above, ending at Imain = 0, and with the trim coils still not powered, measure the length-integrated vertical dipole field, (Bydl, at Imain = 0.  Then open relay to disconnect the power supply from the magnet and measure the length-integrated vertical dipole field again.

	Filename & run number of  (Bydl at Imain = 0:
	Wiredat.ru2, wireplt.ru2

	Filename & run number of  (Bydl when disconnected:
	Wiredat.ru3, wireplt.ru3


11) With the main coils still hooked up, connect the trim coil to a bipolar 6-A supply (MCOR6) with proper trim polarity as determined above.

12) With the trim coil at zero, standardize the magnet as described above in step #8, finishing with the main coil at Imain = 0.  Then measure (Bydl as a function of trim coil current from 0 to 6 A in 0.25-A steps, including zero (25 ‘down’ measurements), and again from 6 to 6 A in 0.5-A steps (25 ‘up’ measurements).  Then set the trim current to zero.

	Filename & run # of (Bydl trim data at Imain = 0:
	Wiredat.ru4, wireplt.ru4


13) Run the best degauss procedure known using cycling of the main coil current (trim current at zero) and record the smallest final measured |(Bydl| achievable and reproducible with Imain = 0.  Please also finish the degauss procedure with a positive step, by setting the current in the positive direction to zero (i.e., from Imain < 0 to Imain = 0).  Record the degauss procedure applied (ramp rate, hold times, current sequence, etc). See file SRXSS Dipole Degauss Procedure.docx for more details.
	De-Gauss procedure’s achieved mean |(Bydl|:
	T-m

	Measurements used for mean and stdev.
	

	Ramp rate:
	2 A/s

	Hold times:
	5 sec

	Current sequence:


	 0.0, 13.0000, -11.4400, 10.0672, -8.8591,  7.7960, -6.8605,  6.0373, -5.3128,  4.6752, -4.1142,  3.6205, -3.1861,  2.8037, -2.4673,  2.1712, -1.9107,  1.6814, -1.4796,  1.3021, -1.1458,  1.0083, -0.8873,  0.7808, -0.6871,  0.6047, -0.5321,  0.4683, -0.4121,  0.3626, -0.3191,  0.2808, -0.2471,  0.2175, -0.1914,  0.1684, -0.1482,  0.1304, -0.1148,  0.1010, -0.0889,  0.0782, -0.0688,  0.0606, -0.0533,  0.0469, -0.0413,  0.0363, -0.0320,  0.0281, -0.0248,  0.0218, -0.0192, 0.0


14) With the degauss procedure finished, the trim coil still at zero, and without having changed the main coil current at all from its Imain = 0 setting after step #13 above, please vary the trim coil current from 0 to +6 A in 0.25-A steps, while measuring the length-integrated vertical dipole field, (Bydl, at each setting, including zero (25 ‘up’ measurements), and again from +6 A to 6 A in 0.5-A steps (25 ‘down’ measurements), and finally from 6 A to +6 A in 0.5-A steps (25 more ‘up’ measurements).  These field integral values will be quite small, so please take care to resolve the measurements at the level of <0.001 kG-m, if possible.
	Filename & run # of (Bydl after degauss:
	Wiredat.ru5, wireplt.ru5

	Poletip Field
	 +/-  G

	Filename & run # of (Bydl vs. trim after degauss:
	Wiredat.ru6, wireplt.ru6


15) For all four dipoles (except as noted below), with stretched wire, and after re-standardization using the bipolar method of step #8, measure the length-integrated vertical field over a horizontal span of (7 mm at each 1-mm interval, at the following main and trim coil current settings.

· Imain = +6 A, and Itrim = 0 (all 4 magnets)
· Imain = +13 A, and Itrim = 0 (all 4 magnets)

	Filename & run # of (Bydl vs. x data at 6 A, 0 A:
	Wirevsx.ru7, wirepltvsx.ru7

	Filename & run # of (Bydl vs. x data at 13 A, 0 A:
	Wirevsx.ru8, wirepltvsx.ru8


16) For the BCXXL1 magnet only, after bipolar standardization (step #8 above), with main coil at +4 A, and with trim at zero, measure the harmonics with a rotating coil with at least a 5-mm diameter but less than 8-mm diameter (if not available we may need to skip this step).  Record probe designation, radius, and data file names:

	Coil designation (text):
	N/A

	Coil radius (mm):
	N/A

	BCXXL1 harmonics filename:
	N/A


17) Measure the inductance and resistance of the main and trim magnet coils and also verify the concurrent magnet temperature:

	Ambient Temp (oC):
	21.7 

	Present magnet temperature on top surface (oC):
	20.3

	Inductance of main coil (mH):
	38.46

	Resistance of main coil (Ohms):
	

	Inductance of trim coil (mH):
	0.064

	Resistance of trim coil (Ohms):
	


18) Measure the pole tip field at 13 A on main coil:

	Pole tip field:
	0.723 T

	Current:
	13.00242 A


19) Upon completion of tests, send traveler to Dorian Bohler and John Amann.  Dorian Bohler will determine if the magnet is accepted.  Upon acceptance of magnet, analysis data will be placed in on-line data folder.

	Magnet accepted and Analysis file(s) put into on-line data folder (initials):
	

	Assigned beamline location (MAD-deck name):
	BCXXL4


[image: image3.jpg]