{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["Sheet1","Sheet2","Sheet3"],"not_rendered":{"charts":0,"images":0,"macros":1,"external_links":0,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"Sheet1","index":0,"n_rows":197,"n_cols":4,"rows_kept":197,"cut":false,"hidden":false,"merged":[]}
["MMF",null,null,"PM Checks"]
[]
["\"X\""]
[]
["Undulator","SN 08"]
["Date:","2013-10-18"]
["Crew:","MJR, LG"]
[]
["BUCK IN "]
["Far Value",0,"PM1 X1"]
["Near Value",0,"PM 4 X1"]
["1st Far",27]
["1st. Near",27.655000000000001]
["2nd. Far",28.5]
["2nd Near",28.03]
["Far Should Read",{"$f":"BuckIn(B10,B11,B12,B13,B14,B15)","v":27.873333333333335}]
["Near Should Read",27.873333333333335]
["LOS",27.873333333333335]
[]
["Undulator "]
[]
[]
["1)","2)","3)","4)"]
[32.369999999999997,32.365000000000002,32.363,32.362000000000002]
[1,1,1,1]
[{"$f":"SUM(A24:A25)","v":33.369999999999997},{"$f":"SUM(B24:B25)","v":33.365000000000002},{"$f":"SUM(C24:C25)","v":33.363},{"$f":"SUM(D24:D25)","v":33.362000000000002}]
[{"$f":"B18","v":27.873333333333335},{"$f":"B18","v":27.873333333333335},{"$f":"B18","v":27.873333333333335},{"$f":"B18","v":27.873333333333335}]
[{"$f":"SUM(A27-A26)","v":-5.4966666666666626},{"$f":"SUM(B27-B26)","v":-5.4916666666666671},{"$f":"SUM(C27-C26)","v":-5.4896666666666647},{"$f":"SUM(D27-D26)","v":-5.488666666666667}]
[]
["5)","6)","7)","8)"]
[22.588000000000001,22.582000000000001,22.582000000000001,22.584]
[1,1,1,1]
[{"$f":"SUM(A31:A32)","v":23.588000000000001},{"$f":"SUM(B31:B32)","v":23.582000000000001},{"$f":"SUM(C31:C32)","v":23.582000000000001},{"$f":"SUM(D31:D32)","v":23.584}]
[{"$f":"B18","v":27.873333333333335},{"$f":"B18","v":27.873333333333335},{"$f":"B18","v":27.873333333333335},{"$f":"B18","v":27.873333333333335}]
[{"$f":"SUM(A34-A33)","v":4.2853333333333339},{"$f":"SUM(B34-B33)","v":4.2913333333333341},{"$f":"SUM(C34-C33)","v":4.2913333333333341},{"$f":"SUM(D34-D33)","v":4.2893333333333352}]
[]
[]
[null,"Reference","Block"]
[]
["A)","B)","C)","D)"]
[33.914000000000001,29.384,27.623999999999999,27.623999999999999]
[1,1,1,1]
[{"$f":"SUM(A41:A42)","v":34.914000000000001},{"$f":"SUM(B41:B42)","v":30.384},{"$f":"SUM(C41:C42)","v":28.623999999999999},{"$f":"SUM(D41:D42)","v":28.623999999999999}]
[{"$f":"B18","v":27.873333333333335},{"$f":"B18","v":27.873333333333335},{"$f":"B18","v":27.873333333333335},{"$f":"B18","v":27.873333333333335}]
[{"$f":"SUM(A44-A43)","v":-7.0406666666666666},{"$f":"SUM(B44-B43)","v":-2.5106666666666655},{"$f":"SUM(C44-C43)","v":-0.75066666666666393},{"$f":"SUM(D44-D43)","v":-0.75066666666666393}]
[]
[]
[]
[]
[]
["MMF",null,null,"PM Checks"]
[]
["\"X\""]
[]
[{"$f":"B18","v":27.873333333333335},"LOS"]
["Undulator",{"$f":"B5","v":"SN 08"}]
["Date:",{"$f":"B6","v":"2013-10-18"}]
["Crew:",{"$f":"B7","v":"MJR, LG"}]
[]
[]
[null,"PM4"]
[]
["X1) Buck In","Y1)","Y2)"]
[null,30.472000000000001,30.475999999999999]
[null,1,1]
[null,{"$f":"SUM(B64:B65)","v":31.472000000000001},{"$f":"SUM(C64:C65)","v":31.475999999999999}]
[null,{"$f":"B18","v":27.873333333333335},{"$f":"B18","v":27.873333333333335}]
[0,{"$f":"SUM(B67-B66)","v":-3.5986666666666665},{"$f":"SUM(C67-C66)","v":-3.6026666666666642}]
[]
[]
[null,"PM3"]
[]
["X1)","Y1)","Y2)"]
[28.765000000000001,32.36,32.36]
[1,1,1]
[{"$f":"SUM(A74:A75)","v":29.765000000000001},{"$f":"SUM(B74:B75)","v":33.36},{"$f":"SUM(C74:C75)","v":33.36}]
[{"$f":"B18","v":27.873333333333335},{"$f":"B18","v":27.873333333333335},{"$f":"B18","v":27.873333333333335}]
[{"$f":"SUM(A77-A76)","v":-1.8916666666666657},{"$f":"SUM(B77-B76)","v":-5.4866666666666646},{"$f":"SUM(C77-C76)","v":-5.4866666666666646}]
[]
[]
[null,"PM 2"]
["X1)","Y1)","Y2)"]
[28.763000000000002,32.356999999999999,32.357999999999997]
[1,1,1]
[{"$f":"SUM(A83:A84)","v":29.763000000000002},{"$f":"SUM(B83:B84)","v":33.356999999999999},{"$f":"SUM(C83:C84)","v":33.357999999999997}]
[{"$f":"B18","v":27.873333333333335},{"$f":"B18","v":27.873333333333335},{"$f":"B18","v":27.873333333333335}]
[{"$f":"SUM(A86-A85)","v":-1.8896666666666668},{"$f":"SUM(B86-B85)","v":-5.4836666666666645},{"$f":"SUM(C86-C85)","v":-5.4846666666666621}]
[]
[]
[null,"PM1"]
["X1) Buck In","Y1)","Y2)"]
[null,30.471,30.466999999999999]
[null,1,1]
[null,{"$f":"SUM(B92:B93)","v":31.471},{"$f":"SUM(C92:C93)","v":31.466999999999999}]
[null,{"$f":"B18","v":27.873333333333335},{"$f":"B18","v":27.873333333333335}]
[0,{"$f":"SUM(B95-B94)","v":-3.5976666666666652},{"$f":"SUM(C95-C94)","v":-3.5936666666666639}]
[]
[]
[]
[]
["MMF",null,null,"PM Checks"]
[]
["\"Y\"",null,"PM4Y1","PM1Y1"]
[null,null,8.9329999999999998,8.9369999999999994]
[null,null,1,1]
["Undulator",{"$f":"B5","v":"SN 08"},{"$f":"SUM(C104+C105)","v":9.9329999999999998},{"$f":"SUM(D104+D105)","v":9.9369999999999994}]
["Date:",{"$f":"B6","v":"2013-10-18"}]
["Crew:",{"$f":"B7","v":"MJR, LG"}]
[null,null,"HI =",{"$f":"AVERAGE(C106,D106)","v":9.9349999999999987}]
[]
["PM4",null,"PM3"]
[]
["Y1)","X1)","Y2)","X1"]
[{"$f":"C104","v":8.9329999999999998},12.538,8.9369999999999994,12.538]
[1,1,1,1]
[{"$f":"SUM(A114:A115)","v":9.9329999999999998},{"$f":"SUM(B114:B115)","v":13.538},{"$f":"SUM(C114:C115)","v":9.9369999999999994},{"$f":"SUM(D114:D115)","v":13.538}]
[{"$f":"D109","v":9.9349999999999987},{"$f":"D109","v":9.9349999999999987},{"$f":"D109","v":9.9349999999999987},{"$f":"D109","v":9.9349999999999987}]
[{"$f":"SUM(A117-A116)","v":1.9999999999988916E-3},{"$f":"SUM(B117-B116)","v":-3.6030000000000015},{"$f":"SUM(C117-C116)","v":-2.0000000000006679E-3},{"$f":"SUM(D117-D116)","v":-3.6030000000000015}]
[]
[]
["PM2",null,"PM1"]
[]
["Y1)","X1)","Y1)","X1)"]
[8.9390000000000001,12.539,{"$f":"D104","v":8.9369999999999994},12.537000000000001]
[1,1,1,1]
[{"$f":"SUM(A124:A125)","v":9.9390000000000001},{"$f":"SUM(B124:B125)","v":13.539},{"$f":"SUM(C124:C125)","v":9.9369999999999994},{"$f":"SUM(D124:D125)","v":13.537000000000001}]
[{"$f":"D109","v":9.9349999999999987},{"$f":"D109","v":9.9349999999999987},{"$f":"D109","v":9.9349999999999987},{"$f":"D109","v":9.9349999999999987}]
[{"$f":"SUM(A127-A126)","v":-4.0000000000013358E-3},{"$f":"SUM(B127-B126)","v":-3.604000000000001},{"$f":"SUM(C127-C126)","v":-2.0000000000006679E-3},{"$f":"SUM(D127-D126)","v":-3.6020000000000021}]
[]
[]
[]
[]
[]
[]
["Top Of Gage Block",null,"Probe Location in (X)"]
[]
[14.333,null,-5.4640000000000004,"LOS"]
[{"$f":"D109","v":9.9349999999999987},null,4.2999999999999997E-2,"Mic To Plus Edge"]
[{"$f":"SUM(A138-A137)","v":-4.3980000000000015},null,{"$f":"SUM(C137+C138)","v":-5.4210000000000003}]
[null,null,-3.9E-2]
[null,null,{"$f":"(C139+C140)","v":-5.46},"Probe Location"]
[]
[]
[]
[]
[]
[]
[]
[]
[]
["MMF",null,null,"PM Checks"]
[]
["\"Y\""]
[]
["Undulator",{"$f":"B5","v":"SN 08"}]
["Date:",{"$f":"B6","v":"2013-10-18"}]
["Crew:",{"$f":"B7","v":"MJR, LG"}]
[]
[]
["HI =",{"$f":"D109","v":9.9349999999999987}]
[]
["Undulator"]
[]
["1)","2)","3)","4)"]
[5.4550000000000001,5.4569999999999999,5.46,5.4560000000000004]
[1,1,1,1]
[{"$f":"SUM(A165:A166)","v":6.4550000000000001},{"$f":"SUM(B165:B166)","v":6.4569999999999999},{"$f":"SUM(C165:C166)","v":6.46},{"$f":"SUM(D165:D166)","v":6.4560000000000004}]
[{"$f":"D109","v":9.9349999999999987},{"$f":"D109","v":9.9349999999999987},{"$f":"D109","v":9.9349999999999987},{"$f":"D109","v":9.9349999999999987}]
[{"$f":"SUM(A168-A167)","v":3.4799999999999986},{"$f":"SUM(B168-B167)","v":3.4779999999999989},{"$f":"SUM(C168-C167)","v":3.4749999999999988},{"$f":"SUM(D168-D167)","v":3.4789999999999983}]
[]
[]
["5)","6)","7)","8)"]
[12.538,12.536,12.535,12.54]
[1,1,1,1]
[{"$f":"SUM(A173+A174)","v":13.538},{"$f":"SUM(B173+B174)","v":13.536},{"$f":"SUM(C173+C174)","v":13.535},{"$f":"SUM(D173+D174)","v":13.54}]
[{"$f":"D109","v":9.9349999999999987},{"$f":"D109","v":9.9349999999999987},{"$f":"D109","v":9.9349999999999987},{"$f":"D109","v":9.9349999999999987}]
[{"$f":"SUM(A176-A175)","v":-3.6030000000000015},{"$f":"SUM(B176-B175)","v":-3.6010000000000009},{"$f":"SUM(C176-C175)","v":-3.6000000000000014},{"$f":"SUM(D176-D175)","v":-3.6050000000000004}]
[]
[]
[]
["Reference Block"]
[]
["A)","B)","C)","D)"]
[9.1739999999999995,null,11.3118,14.4]
[1,null,1,1]
[{"$f":"SUM(A184+A185)","v":10.173999999999999},null,{"$f":"SUM(C184+C185)","v":12.3118},{"$f":"SUM(D184+D185)","v":15.4}]
[{"$f":"D109","v":9.9349999999999987},null,{"$f":"D109","v":9.9349999999999987},{"$f":"D109","v":9.9349999999999987}]
[{"$f":"SUM(A187-A186)","v":-0.23900000000000077},0,{"$f":"SUM(C187-C186)","v":-2.3768000000000011},{"$f":"SUM(D187-D186)","v":-5.4650000000000016}]
[]
[]
["Flats"]
[]
["U/S",null,"D/S"]
["(+)","(-)","(+)","(-)"]
[8.2189999999999994,8.2200000000000006,8.2200000000000006,8.2210000000000001]
[{"$f":"D109","v":9.9349999999999987},{"$f":"D109","v":9.9349999999999987},{"$f":"D109","v":9.9349999999999987},{"$f":"D109","v":9.9349999999999987}]
[{"$f":"SUM(A196-A195)","v":1.7159999999999993},{"$f":"SUM(B196-B195)","v":1.7149999999999981},{"$f":"SUM(C196-C195)","v":1.7149999999999981},{"$f":"SUM(D196-D195)","v":1.7139999999999986}]

{"sheet":"Sheet2","index":1,"n_rows":42,"n_cols":8,"rows_kept":42,"cut":false,"hidden":false,"merged":[]}
[]
["Undulator"," ",{"$f":"Sheet1!B5","v":"SN 08"}]
["Date"," ",{"$f":"Sheet1!B6","v":"2013-10-18"}]
["Crew"," ",{"$f":"Sheet1!B7","v":"MJR, LG"}]
[]
[]
["Undulator Values"]
[]
["TB","Y (in)","Y (mm)",null,null,"TB","X (in)","X (mm)"]
[1,{"$f":"Sheet1!A169","v":3.4799999999999986},{"$f":"B10*25.4","v":88.391999999999967},null,null,5,{"$f":"Sheet1!A35","v":4.2853333333333339},{"$f":"G10*25.4","v":108.84746666666668}]
[2,{"$f":"Sheet1!B169","v":3.4779999999999989},{"$f":"B11*25.4","v":88.341199999999972},null,null,6,{"$f":"Sheet1!B35","v":4.2913333333333341},{"$f":"G11*25.4","v":108.99986666666668}]
[3,{"$f":"Sheet1!C169","v":3.4749999999999988},{"$f":"B12*25.4","v":88.264999999999958},null,null,7,{"$f":"Sheet1!C35","v":4.2913333333333341},{"$f":"G12*25.4","v":108.99986666666668}]
[4,{"$f":"Sheet1!D169","v":3.4789999999999983},{"$f":"B13*25.4","v":88.366599999999949},null,null,8,{"$f":"Sheet1!D35","v":4.2893333333333352},{"$f":"G13*25.4","v":108.94906666666671}]
[]
[]
["PM Values"]
[]
[null,null,"Upstream"," ","(PM 3)",null,"Checks"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)",null,"X (mm)","Y (mm)"]
["Y2",{"$f":"Sheet1!C78","v":-5.4866666666666646},{"$f":"Sheet1!C118","v":-2.0000000000006679E-3},{"$f":"B20*25.4","v":-139.36133333333328},{"$f":"C20*25.4","v":-5.0800000000016964E-2},null,{"$f":"H10-D21","v":156.89579999999998},{"$f":"C10-E20","v":88.442799999999991}]
["X1",{"$f":"Sheet1!A78","v":-1.8916666666666657},{"$f":"Sheet1!D118","v":-3.6030000000000015},{"$f":"B21*25.4","v":-48.048333333333304},{"$f":"C21*25.4","v":-91.51620000000004}]
["Y1",{"$f":"Sheet1!B78","v":-5.4866666666666646},null,{"$f":"B22*25.4","v":-139.36133333333328}]
[]
[null,null,"Downstream",null,"(PM 2)"," ","Checks"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)",null,"X (mm)","Y (mm)"]
["Y1",{"$f":"Sheet1!B87","v":-5.4836666666666645},{"$f":"Sheet1!A128","v":-4.0000000000013358E-3},{"$f":"B26*25.4","v":-139.28513333333328},{"$f":"C26*25.4","v":-0.10160000000003393},null,{"$f":"H13-D27","v":156.94660000000005},{"$f":"C13-E26","v":88.468199999999982}]
["X1",{"$f":"Sheet1!A87","v":-1.8896666666666668},{"$f":"Sheet1!B128","v":-3.604000000000001},{"$f":"B27*25.4","v":-47.997533333333337},{"$f":"C27*25.4","v":-91.541600000000017}]
["Y2",{"$f":"Sheet1!C87","v":-5.4846666666666621},"\\",{"$f":"B28*25.4","v":-139.31053333333321}]
[]
[]
[]
[]
[null,"Probe",null,null,null,"Flats"]
[]
[null,"X (in)","X (mm)",null," ","Y(in)","Y(mm)"]
[null,{"$f":"Sheet1!C141","v":-5.46},{"$f":"B36*25.4","v":-138.684},null,"US  +",{"$f":"Sheet1!A197","v":1.7159999999999993},{"$f":"F36*25.4","v":43.586399999999983}]
[null," ",null,null,"US   -",{"$f":"Sheet1!B197","v":1.7149999999999981},{"$f":"F37*25.4","v":43.56099999999995}]
[null,null,null,null,"DS  +",{"$f":"Sheet1!C197","v":1.7149999999999981},{"$f":"F38*25.4","v":43.56099999999995}]
[null,null,"Top of Gage Block",null,"DS   -",{"$f":"Sheet1!D197","v":1.7139999999999986},{"$f":"F39*25.4","v":43.53559999999996}]
[]
[null,"Y (in)","Y (mm)"]
[null,{"$f":"Sheet1!A139","v":-4.3980000000000015},{"$f":"B42*25.4","v":-111.70920000000002}]

{"sheet":"Sheet3","index":2,"n_rows":52,"n_cols":9,"rows_kept":52,"cut":false,"hidden":false,"merged":[]}
[]
["Undulator"," ",13]
["Date"," ",{"$f":"Sheet1!B6","v":"2013-10-18"}]
["Crew"," ","MR,BR"]
[]
[]
[null,null,"Undulator Values"]
[]
["TB","X (in)","Y (in)","X (mm)","Y (mm)",null,"Probe"]
[1,{"$f":"Sheet1!A28","v":-5.4966666666666626},{"$f":"Sheet1!A169","v":3.4799999999999986},{"$f":"SUM(B10*25.4)","v":-139.61533333333321},{"$f":"SUM(C10*25.4)","v":88.391999999999967}]
[2,{"$f":"Sheet1!B28","v":-5.4916666666666671},{"$f":"Sheet1!B169","v":3.4779999999999989},{"$f":"SUM(B11*25.4)","v":-139.48833333333334},{"$f":"SUM(C11*25.4)","v":88.341199999999972},null,"X (in)","X (mm)"]
[3,{"$f":"Sheet1!C28","v":-5.4896666666666647},{"$f":"Sheet1!C169","v":3.4749999999999988},-139.43753333333328,88.264999999999958,null,{"$f":"Sheet1!C141","v":-5.46},{"$f":"SUM(G12*25.4)","v":-138.684}]
[4,{"$f":"Sheet1!D28","v":-5.488666666666667},{"$f":"Sheet1!D169","v":3.4789999999999983},-139.41213333333334,88.366599999999949]
[5,{"$f":"Sheet1!A35","v":4.2853333333333339},{"$f":"Sheet1!A177","v":-3.6030000000000015},108.84746666666668,-91.51620000000004]
[6,{"$f":"Sheet1!B35","v":4.2913333333333341},{"$f":"Sheet1!B177","v":-3.6010000000000009},108.99986666666668,-91.465400000000017,null,"Top of Gage Block"]
[7,{"$f":"Sheet1!C35","v":4.2913333333333341},{"$f":"Sheet1!C177","v":-3.6000000000000014},108.99986666666668,-91.440000000000026]
[8,{"$f":"Sheet1!D35","v":4.2893333333333352},{"$f":"Sheet1!D177","v":-3.6050000000000004},108.94906666666671,-91.567000000000007,null,"Y (in)","Y (mm)"]
[null,null,null,null,null,null,{"$f":"Sheet1!A139","v":-4.3980000000000015},{"$f":"SUM(G18*25.4)","v":-111.70920000000002}]
[null,null,"Undulator PM Values"]
[]
[null,null,"Upstream"," ","(PM 3)"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)",null,"Flats"]
["Y2",{"$f":"Sheet1!A78","v":-1.8916666666666657},{"$f":"Sheet1!C118","v":-2.0000000000006679E-3},{"$f":"SUM(B23*25.4)","v":-48.048333333333304},{"$f":"SUM(C23*25.4)","v":-5.0800000000016964E-2}]
["X1",{"$f":"Sheet1!A78","v":-1.8916666666666657},{"$f":"Sheet1!D118","v":-3.6030000000000015},{"$f":"SUM(B24*25.4)","v":-48.048333333333304},{"$f":"SUM(C24*25.4)","v":-91.51620000000004},null," ","Y(in)","Y(mm)"]
["Y1",{"$f":"Sheet1!B78","v":-5.4866666666666646},null,{"$f":"SUM(B25*25.4)","v":-139.36133333333328},null,null,"US  +",{"$f":"Sheet1!A197","v":1.7159999999999993},{"$f":"SUM(H25*25.4)","v":43.586399999999983}]
[null,null,null,null,null,null,"US   -",{"$f":"Sheet1!B197","v":1.7149999999999981},{"$f":"SUM(H26*25.4)","v":43.56099999999995}]
[null,null,"Downstream",null,"(PM 2)",null,"DS  +",{"$f":"Sheet1!C197","v":1.7149999999999981},{"$f":"SUM(H27*25.4)","v":43.56099999999995}]
["TB","X (in)","Y (in)","X (mm)","Y (mm)",null,"DS   -",{"$f":"Sheet1!D197","v":1.7139999999999986},{"$f":"SUM(H28*25.4)","v":43.53559999999996}]
["Y1",{"$f":"Sheet1!B87","v":-5.4836666666666645},{"$f":"Sheet1!A128","v":-4.0000000000013358E-3},{"$f":"SUM(B29*25.4)","v":-139.28513333333328},{"$f":"SUM(C29*25.4)","v":-0.10160000000003393}]
["X1",{"$f":"Sheet1!A87","v":-1.8896666666666668},{"$f":"Sheet1!B128","v":-3.604000000000001},{"$f":"SUM(B30*25.4)","v":-47.997533333333337},{"$f":"SUM(C30*25.4)","v":-91.541600000000017}]
["Y2",{"$f":"Sheet1!C87","v":-5.4846666666666621},null,{"$f":"SUM(B31*25.4)","v":-139.31053333333321}]
[]
[null,null,"Ref. PM Values"]
[]
[null,null,"Upstream"," ","(PM 4)"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)"]
["Y1",{"$f":"Sheet1!B68","v":-3.5986666666666665},{"$f":"Sheet1!A118","v":1.9999999999988916E-3},{"$f":"SUM(B37*25.4)","v":-91.406133333333329},{"$f":"SUM(C37*25.4)","v":5.079999999997184E-2},null,null," "," "]
["X1",0,{"$f":"Sheet1!B118","v":-3.6030000000000015},0,{"$f":"SUM(C38*25.4)","v":-91.51620000000004}]
["Y2",{"$f":"Sheet1!C68","v":-3.6026666666666642},null,{"$f":"SUM(B39*25.4)","v":-91.507733333333263}]
[]
[null,null,"Downstream",null,"(PM 1)"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)"]
["Y1",{"$f":"Sheet1!B96","v":-3.5976666666666652},{"$f":"Sheet1!C128","v":-2.0000000000006679E-3},{"$f":"SUM(B43*25.4)","v":-91.380733333333296},{"$f":"SUM(C43*25.4)","v":-5.0800000000016964E-2}]
["X1",0,{"$f":"Sheet1!D128","v":-3.6020000000000021},0,{"$f":"SUM(C44*25.4)","v":-91.49080000000005}]
["Y2",{"$f":"Sheet1!C96","v":-3.5936666666666639},null,{"$f":"SUM(B45*25.4)","v":-91.279133333333263}]
[]
[null,null,"Ref. Block"]
["TB","X (in)","Y (in)","X (mm)","Y (mm)"]
["A",{"$f":"Sheet1!A45","v":-7.0406666666666666},{"$f":"Sheet1!A188","v":-0.23900000000000077},{"$f":"SUM(B49*25.4)","v":-178.83293333333333},{"$f":"SUM(C49*25.4)","v":-6.0706000000000193}]
["B",{"$f":"Sheet1!B45","v":-2.5106666666666655},null,{"$f":"SUM(B50*25.4)","v":-63.770933333333296}]
["C",{"$f":"Sheet1!C45","v":-0.75066666666666393},{"$f":"Sheet1!C188","v":-2.3768000000000011},{"$f":"SUM(B51*25.4)","v":-19.066933333333264},{"$f":"SUM(C51*25.4)","v":-60.370720000000027}]
["D",{"$f":"Sheet1!D45","v":-0.75066666666666393},{"$f":"Sheet1!D188","v":-5.4650000000000016},{"$f":"SUM(B52*25.4)","v":-19.066933333333264},{"$f":"SUM(C52*25.4)","v":-138.81100000000004}]