Additional EPU measurements

- Measure Bx and By reference magnets’ field integrals using: 1) the long coil at various X positions of the coil 2) the flip coil in bldg. 26 3) the taught wire in bldg. 26 4) Hall probe scans; compare results

A Bx reference magnet was measured first in Bldg.26 lab by a 0.4m long stretched wire. At MMF we measured the same magnet with the coil at different x positions. To set the current we used HP E3610A Power Supply with Danfysik Ultrastab 860R current transducer and Danfysik 860S range selector. The magnet was standardized by cycling the current 2 times; 0A – 5A – 0A. The current set to 0.998A. Undulator gap set to 113mm. Magnetic axis is at x=0.

Results are in the following table (‘Integral’ is ‘Coil’ – ‘Current=0’).

| X (mm) | Current = 0 | Coil (G-cm) | Integral (G-cm) | Wire (G-cm) | Diff. (G-cm) | Diff. (%) |
| --- | --- | --- | --- | --- | --- | --- |
| -24 | +87 | -396 | -483 | -476 | -7 | -1.5 |
| 0 | +104 | -371 | -475 | -476 | +1 | +0.2 |
| +24 | +120 | -361 | -481 | -476 | -5 | -1.3 |

Hall probe measurement: I1X = -481 G-cm (1.2%).

- Measure EPU in VP- at 13 mm and 30 mm gap using the long coil and the Hall probe

Coil data in Run 42 and Run 43 at: v:\MET\MagServe\BL5_EPU\Z Scans\.

- In CP+ and CP- polarities, 13 mm gap, shift all rows by  Z = +1/8 * period and by  Z =  -1/8 * period and measure with the long coil; compare to unshifted CP + and CP- measurements

The same skew quadrupole.

Data in Run 46, 47, 48, 49 at:  v:\MET\MagServe\BL5_EPU\Long Coil Scans\

- Perform scans off the magnetic plane  in CP and VP and confirm integrated field measurements obey Maxwell’s equations (divergence and curl)

Theory:

Let Fx = Bxdz and Fy = Bydz

Then ·B = ∂Fx∂x+ ∂Fy∂y=0   and ×B = ∂Fy∂x- ∂Fx∂y=0

Measurements:

To calculate derivatives we have measured field integrals vs. x at y=0 and field integrals vs. y at x = 0 using long coil. No Earth field corrections.

Coil measurement results are in the following table (In G-cm):

|  |  | CP+ |  | CP- |  | VP+ |  | VP- |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Y(mm) | X(mm) | I1Y | I1X | I1Y | I1X | I1Y | I1X | I1Y | I1X |
| 0 | -2 | -248 | +205 | -98 | -59 | -589 | +387 | -339 | -226 |
| 0 | 0 | +169 | +42 | +307 | +94 | +510 | +188 | +641 | -47 |
| 0 | +2 | +526 | -106 | +602 | +244 | +1387 | -40 | +1339 | +193 |
| -2 | 0 | +21 | -434 | +546 | -311 | +479 | -1086 | +1083 | -1074 |
| 0 | 0 | +183 | +17 | +340 | +86 | +559 | -30 | +682 | -72 |
| +2 | 0 | +380 | +452 | +214 | +495 | +836 | +985 | +599 | +863 |
| Derivatives(G) |  |  |  |  |  |  |  |  |  |
| /dx |  | +1935 | -778 | +1750 | +758 | +4940 | -1068 | +4195 | +1048 |
| /dy |  | +898 | +2215 | -830 | +2015 | +892 | +5178 | -1210 | +4842 |
| ·B |  | +120 |  | -72 |  | -176 |  | -162 |  |
| ×B |  | -280 |  | -265 |  | -238 |  | -647 |  |
| div./ curl (% of slope) |  | 15%/14% |  | 9%/13% |  | 20%/5% |  | 15%/14% |  |

Maxwell's equations are satisfied with a reasonable accuracy.

Hall probe measurement results are in: s:\MagData\LCLS\Undulator\SSRL-BL5-EPU\DATASET0001\Rough Tuning\Z Scans\Run 033÷037

Summary is in the table (Integrals in G-cm). Measurements are corrected for Earth field by adding constants:  X = -0.2G; Y = +0.4G

|  |  | CP- |  |
| --- | --- | --- | --- |
| Y(mm) | X(mm) | I1Y | I1X |
| 0 | 0 | +538 | -72 |
| 0 | -1 | +828 | +26 |
| 0 | +1 | +202 | -174 |
| -0.5 | 0 | +578 | +75 |
| +0.5 | 0 | +501 | -224 |
| Derivatives(G) |  |  |  |
| /dx |  | -3130 | -1000 |
| /dy |  | -770 | -3000 |
| ·B |  | -1770 ?!! |  |
| ×B |  | -130 OK! |  |
| div./ curl (% of slope) |  | 170%/4% |  |