[image: image1.jpg]I / . I C' Linac Coherent Light Source Stanford Linear Accelerator Center
e —— e = 2 =
LA

Stanford Synchrotron Rediation Laboratory




[image: image6.png]





[image: image7.jpg]



	LCLS ‘Q150kG’ LTU QUADRUPOLE MAGNET

	FIDUCIALIZATION REPORT


	




	Inspector:
	Keith Caban

	Responsible Engineer:
	Carl Rago

	Date:
	Wednesday, December 05, 2007

	Work Order/Charge No.:
	9242609

	Serial Number:
	SLAC – 002013 / SN 014

	URL of Fiducial Report:
	\\Web002\www-group\met\Quality\FIDUCIAL REPORTS\LCLS LTU Q150kG QUADS\002013.pdf


	Part Set-up – Coordinate System Set-up


	Spatial Alignment


· Geometric axis of the poles of the magnet.
	Planar Alignment


· Plane on top of magnet where tooling ball sockets are welded to.
	“Z” Zero


· Mid-plane of the magnet (middle of upstream and downstream ends).
	“X” & “Y” Zero


· Geometric axis of the poles of the magnet.
	[image: image3.jpg]



Front View
	


Side View


	Tooling Ball Measurements/Locations


[image: image5.jpg]3, overson tesla
* o’ incorporated
/& _[@~

LCLS LTU Quadrupole Magnet
Serial #: 001

SLAC Pog# 62736

Magnet Weight: 6374

Passed Mechanical

Passed Elecirical

_Made In U.S.A.

Mg g




Tooling Ball Adapter Cylinder Projected 1” Offset to the Tooling Ball Adapter Plane
	Tooling Ball
	Cyl. Dia.
	X
	Y
	Z

	Proj. TB 1
	0.25005
	2.66651
	10.21003
	5.60837

	Proj. TB 2
	0.24992
	2.66788
	10.20399
	-5.60948

	Proj. TB 3
	0.25009
	-2.64442
	10.20704
	-5.61091

	Proj. TB 4
	0.24999
	-2.64199
	10.21256
	5.60608


Pole Data

*Data looking from Downstream End
	Side
	Pole Dist A-C
	Pole Dist B-D
	Gap 1
	Gap 2
	Gap 3
	Gap 4

	Downstream End (+Z)
	1.26063
	1.26113
	0.43383
	0.43667
	0.43336
	0.43231

	Upstream End (-Z)
	1.25891
	1.25888
	0.43302
	0.43345
	0.43074
	0.43088






























+Y





+Z








+Y





+X











+Z





+X




















