[image: image4.png]



[image: image5.png]





	LCLS Injector Quadrupole Magnet

	FIDUCIALIZATION REPORT


	[image: image1.jpg]





	Inspector:
	Keith Caban

	Responsible Engineer:
	Roger Carr

	Date:
	Monday, October 03, 2005

	Work Order/Charge No.:
	N/A

	Serial Number
	000408


	Part Set-up – Coordinate System Set-up


	Planar Alignment


· Mid-Plane of the magnet
	Spatial Alignment


· A line on the top part of the magnet
· +X goes towards (Magnet Info Label)
	“Z” Zero


· Mid-Plane of the magnet
	“X” & “Y” Zero


· On both ends
· Tangent point of each radii (4 on each end, 8 total).
· Create a line between diagonal tangent points creates 2 lines.
· Intersect the lines.
· Creates a point on each end.
· Create a line of these 2 end points
· This is the “X” & “Y” Zero, and Beamline or” Z” Axis.  
[image: image2.jpg]



	Tooling Ball Measurements/Locations


Top of magnet; view from “+Y”





	Tooling Ball
	FORM
	DIAMETER
	X
	Y
	Z

	TB 1
	0.00077
	0.49549
	6.50043
	8.87843
	-1.24861

	TB 2
	0.00045
	0.49802
	-6.49959
	8.87643
	-1.24936

	TB 3
	0.00049
	0.49688
	-6.50041
	8.87714
	1.25042

	TB 4
	0.00066
	0.49616
	6.50140
	8.87805
	1.25143


Additional Requested Measurements


Tangent Point Straightness

A. 0.00025 

B. 0.00021

C. 0.00032

D. 0.00049

Parallelism to Beamline 

A. 0.00114

B. 0.00064

C. 0.00129

D. 0.00049

Distance of opposite Tangent Axis


A-C = 1.26016 (0.63009, 0.63007)

B-D = 1.25877 (0.62945, 0.62932)

[image: image3.emf]Profile of Pole Comparisons

-0.435

-0.415

-0.395

-0.375

-0.355

-0.335

-0.315

-0.295

-0.275

-0.255

-0.760 -0.710 -0.660 -0.610 -0.560 -0.510 -0.460

X

Y

Scan A

Scan B

Scan C

Scan D


TB 1





TB 2





TB 3





TB 4





Z+





+Z





+X





A





B





C





D








View From +Z





4 Scans, Trasformed and compared





The most maximum deviation  x,y location  (-.750,-.260) is 0.0011” (1.1 mils)  from all 4 scans (green to cyan). 





Approximate Tangent Point 


Area








