## Slide 1
Tutorial On Fiducialization Of Accelerator Magnets And Undulators
Zachary Wolf
SLAC
IMMW19
October 25, 2015

## Slide 2: Overview
Fiducialization Overview
Dipoles
CMM
Capacitive sensors
Moving wire
Quadrupoles, Sextupoles, etc.
CMM
Rotating coil
Vibrating wire
How to locate a wire
Pulsed wire
Solenoids
Small rotating coil
Undulators
CMM
Capacitive sensors
Hall probe measurements
How to locate a Hall probe

## Slide 3: Fiducialization
Suppose we know the beam line coordinates Xb and Yb
in the telescope’s coordinate system.
We know Δx1, Δx2, Δy1, and Δy2 from the fiducialization.
We put the tooling balls at Xb+ Δx1, Xb+ Δx2, Yb+ Δy1, Yb+ Δy2
in order to place the magnet on the beam line.
In this way, we don’t need to know the details of the magnet’s
magnetic field in the tunnel.
Fiducialization: Find the tooling ball locations relative to the
beam axis.

## Slide 4: Dipoles Mechanical Fiducialization
Steel dipoles can be fiducialized mechanically.
A CMM can measure the position of the poles
and relate the pole positions to the tooling ball
positions.
In steel magnets, the magnetic field shape is
‘determined’ by the steel.
Typically the CMM can only touch the poles near the ends and the middle
is ignored. Adequate for short magnets. Also, the poles have some roughness.
Roll accuracy estimate: ~25 µm / 5 cm = 5 x 10-4 rad

## Slide 5: Dipoles Mechanical Fiducialization
Capacitive sensors can measure a reference pole
that has been fiducialized, and then measure the
magnet. Tooling balls on the magnet can be related
to tooling balls on the reference pole.
Sensors
Reference Pole
Electronic Level
Position Detector
Measures over entire length of C-magnets. Need to correct for
sensor position and roll.
Roll accuracy estimate: ~10 µm / 10 cm = 1 x 10-4 rad

## Slide 6: Dipoles Magnetic Fiducialization, Moving Wire
Move wire
vertically,
measure ΔΦ.
Rotate magnet
until ΔΦ=0.
Move wire
horizontally,
measure ΔΦ.
Move magnet
until ΔΦ profile
is centered on
wire.
Similar for finding
vertical center.
Relate magnet position to wire.
Roll
X
Roll accuracy estimate: ~10 µTm / 1 Tm = 1 x 10-5 rad

## Slide 7: Field Integrals From A Moving Wire
V+ = dΦ/dt + Vt
V- = -dΦ/dt + Vt
VT = (VT+ - VT-) / 2
= ΔΦ
Take out thermal emf:

## Slide 8: Quadrupoles And Higher Multipoles Mechanical Fiducialization
CMM probe touches poles
to measure point locations
on the poles.
An algorithm finds the mechanical
center from the sample points.
The mechanical center is related to
tooling balls on the magnet.
Touch poles at quadrupole ends. The middle is ignored. The poles have a roughness.
Quadrupole center position accuracy estimate: ~50 µm.

## Slide 9: Quadrupole Fiducialization Rotating Coil
The dipole strength and the quadrupole strength can be used to
calculate the measurement position relative to the magnetic center.
Quadrupole:

## Slide 10: Quadrupoles And Higher Multipoles Rotating Coil
Locate axis of rotation from shaft ends.
Locate tooling balls relative to axis of rotation.
Harmonic feed down gives magnetic center
location relative to axis of rotation.
From these, calculate tooling ball location
relative to magnetic center.
Can also use short coils.
Move coil until dipole signal
goes to zero.

## Slide 11: Rotating Coil Analysis Field In A Multipole Magnet
R = ‘reference radius’
Assume ∂t = 0, ∂z = 0
θspn = angle of first south pole
of the n’th harmonic

## Slide 12: Voltage From A Rotating Coil
Main harmonic

## Slide 13: Translations Of The Coil Relative To the Magnet Center Are Best Studied In Cartesian Coordinates
We could have started here
since B* is analytic, but we
would have lost the connection
to the field strength and south
pole angle.

## Slide 14: Magnetic Center Relative To The Coil Rotation Axis
Harmonics measured by the coil

## Slide 15: Example: Harmonics Measurement Of A Quadrupole

## Slide 16: Moving Wire Fiducialization
Move the wire to the location where ∫By dz = 0 and ∫Bx dz = 0.
Locate the wire relative to the tooling balls.
(May need to take out pitch and yaw for some applications.)
Quadrupole:
Bx = G y
By = G x
Quadrupole center position accuracy estimate: ~10 µTm / 100 T = small.
Accuracy dominated by locating wire and locating tooling balls.

## Slide 17: Pulsed Wire, Vibrating Wire
The wire has a force on it and moves when a current flows through
the wire in the presence of a magnetic field.
*

## Slide 18: Quadrupoles And Higher Multipoles Vibrating Wire

## Slide 19: Vibrating Wire
If the ac current carrying wire goes
through a magnetic field, the force
excites a mechanical vibration of
the wire.
An optical detector gives
a signal that the wire is vibrating.
The vibration stops when the wire
is at the center of the quadrupole.
This method can also detect pitch and
yaw of the quadrupole.

## Slide 20: Vibrating Wire
Equation of motion of the wire
Solution for the wire motion
Vibration at ωn goes to zero
when Bxn goes to zero.

## Slide 21: Wire Vibration Detector
Optical sensors
detect wire vibration.

## Slide 22: Vibrating Wire Signals
Examples:
Amplitude of signal from wire vibration sensor vs quadrupole x-position
and quadrupole yaw.

## Slide 23: How To Locate A Wire Touching A Wire Moves The Wire
Example:
T = 9 N
L = 2 m
Δx = 50 µm
F = 9 x 10-4 N
Need a better method.

## Slide 24: How To Locate A Wire
Find the center of the wire by
finding its edges.
Calibrate the scale offset by
flipping the detector in a
kinematic mount.

## Slide 25: Wire Sag
The wire position is not at
the end position because
of sag.
There are two ways to deal
with this:
Measure the wire position
near the magnet. The
previous slide illustrated this.
2) Correct for sag.
Sag Correction
Measure at different tensions.
Extrapolate to infinite tension.
At infinite tension, the stage
position gives the wire position
at the magnet.
(earlier slide)

## Slide 26: Pulsed Wire Measurements
Pulsed wire example:
Signal from wire motion of a
pulsed wire in an undulator.
Calibration magnet puts in
step at the end.
When the wave hits the
end of the wire, it reflects.
S. Anderson
In a multipole magnet,
move the magnet or the
wire until there is no
field along the wire, so there
is no signal from the wire
motion detector.
Fiducialize the magnet so
the beam axis is along the
wire.

## Slide 27: Pulsed Wire
Initial impulse
Travelling wave
Derivative of wire position
signal gives field

## Slide 28: Solenoid Fiducialization Pitch and Yaw
A small rotating coil in the center of
the solenoid can be used to align the
field to the axis of rotation.

## Slide 29: Solenoid Fiducialization X and Y
At the end of the solenoid, the rotating coil gives
zero voltage out when the axis of rotation is on
the magnetic center line.

## Slide 30: Undulator Fiducialization Mechanical Measurements
Capacitive sensors can be used to mechanically align a hybrid undulator
to a reference pole. Alternatively, a CMM can be used.

## Slide 31: Undulator Fiducialization Hall Probe Measurements
Delta undulator quadrant
LCLS undulator

## Slide 32: Undulator Fiducialization Tolerance Estimate
This is the tolerance installed in the
tunnel. The fiducialization tolerance
is smaller, around 40 microns.

## Slide 33: Undulator Magnetic Alignment
After undulator is tuned:
X trajectory good.
Y trajectory good.
Phase errors good.
Want the beam to go where
the Hall probe is measuring.
But where is the Hall probe
actually measuring?
Fiducialization tolerance
40 microns.
Probe position given to
100 microns in specs.

## Slide 34: Where is the Hall probe measuring?
Attach to undulator ends, measure offsets
Pointed magnets give a magnetic location that can be found
repeatably at the 2 micron level.

## Slide 35: Pointed Magnet Calibration
Flip magnet in fixture. Distance zero field
point moves gives zero field position
relative to tooling balls.
Yurii Levashov

## Slide 36: Undulator Fiducialization
Want D
D1 : Magnetic Measurement
D2 : Calibration
D3 : CMM
D = D1 + D2 + D3

## Slide 37: Undulator Fiducialization

## Slide 38: Check Fiducialization With Simple Techniques
Check fiducialization.
Use telescopes, pointed magnets,
optical sensor to locate Hall probe

## Slide 39: End