## Slide 1: LCLS-II Undulator Measurements
Zachary Wolf
XFEL
January, 2017

## Slide 2: Outline
  Delta-II
    Position dependence of the K value
    Assembly tolerance to minimize phase errors
    Measurement/assembly plan on CMM
  LCLS-II
    Undulator tuning plan
    K vs gap fits
    Phase matching
    Radiation damage plans
    Earth field strategy
    Phase shifter measurements

## Slide 3: Delta-II K Value Position Dependence
Swiss Light Source

## Slide 4: Delta: Horizontal and Vertical Position Sensitivity LCLS MD 11/17/2015
m-1
µm
mJ
µm
mJ
mJ
µm
µm
/ = 0.942
LCLS MD 11/17/2015
LCLS MD 11/17/2015
[compare to m-1 from model]*
*Alexander Temnikh, private communication
Heinz-Dieter Nuhn

## Slide 5: K Position Dependence Effect On Delta-II
Space for 2-stage HXRSS
Space for 3 Polarization Control Undulators
Location for upgraded HXRSS
Space for SXRSS
What causes the asymmetric position dependence of K?
There is only hyperbolic cosine position dependence in the model of two crossed adjustable phase planar undulators.
What are the alignment tolerance effects on Delta-II?
Is there a model to help explain this behavior so that we can make predictions for Delta-II?
Are there design changes that can minimize this effect?

## Slide 6: A Model To Explain The Delta-II K Value Position Dependence
Add field rolloff to the model of two crossed, planar, adjustable phase undulators.
LCLS-TN-16-1
For one quadrant
Model to explain the position dependence of K
Potential decreases exponentially away from the quadrant, decreases away
from the center line, varies sinusoidally in z, obeys Laplace’s equation.
This potential can be considered the dominant term in a Fourier expansion.
Delta Measurements:
ku = 196 1/m
kr = 270 1/m
ks = 186 1/m

## Slide 7: Delta-II K Value Position Dependence
Lab Frame
Undulator Frame
Do calculations in the undulator frame.
Potential for each quadrant:

## Slide 8: Delta-II K Value Position Dependence
Two crossed adjustable phase undulators.
Combine quadrants 1 and 3, also 2 and 4.

## Slide 9: Delta-II K Value Position Dependence
Scalar potential for all row settings:
Δ13 and Δ24 give the shift between the rows of the adjustable phase undulators. This sets their strength.
Z13 and Z24 set the average position of each adjustable phase undulator. This sets the relative phase between them.

## Slide 10: Delta-II K Value Position Dependence
The fields are calculated from the scalar potential. B = grad Φ

## Slide 11: Delta-II K Value Position Dependence
Knowing the fields, the velocity
components can be found:
Knowing the velocity,
the slippage can be found:
The K value is calculated from the constant part of dS/dz:

## Slide 12: Delta-II K Value Position Dependence
This is the K value to second order in position away from the magnetic
center line (x=0, y=0). (x0, y0) is the beam position.
Note that there is a linear term. This causes alignment problems.

## Slide 13: Delta-II K Value Position Dependence
LCLS-TN-16-1
In the linear modes K vs x does not have a slope at the beam axis. In the circular modes, however, K vs x does have a slope at the beam axis.
This can make setting several undulators to the same K value almost impossible using calibration and alignment techniques.
Undulator frame.
Full formula for K
Approximate K
near center
Linear mode
Circular mode

## Slide 14: Delta-II K Value Position Dependence
Going to the laboratory frame and setting the relative shifts of the two
variable phase undulators to be the same, the formula simplifies.
There is a gradient in x, but not in y.
LCLS-TN-16-1
Note that the gradient comes from the field rolloff ks. It grows as the row shift Δ increases. It is largest in the circular modes where δ = π/2.

## Slide 15: With A Delta-like Undulator Predicted DELTA-II RMS X-Resonance Width
Heinz-Dieter Nuhn
No gap change.
Change K by shifting rows.

## Slide 16: Delta-II K Value Position Dependence
The gradient depends on ks. It is zero for crossed planar undulators.
The gradient gets bigger as the row shift gets bigger, or as the K value gets smaller.
The gradient is zero in the linear modes, and is largest in the circular modes.
The gradient will cause alignment problems with the three Delta-II undulators.
Remedy:
Operate at small relative row shift and increase the gap to adjust K.
Franz Peters’ design
Tried to shape magnets to minimize ks, but no success. Limited by block size.

## Slide 17: Delta-II Construction Tolerances
By is increasing with Z
Delta Measured Peak Fields, LPVMF mode, Krel = 0.9

## Slide 18: Delta-II Construction Tolerances
Delta Measured Phase Errors, LPVMF mode, Krel = 0.9

## Slide 19: Simulation Phase Error With Undulator Taper
LCLS-TN-15-2

## Slide 20: Undulator Taper Conclusion
LCLS-TN-15-2

## Slide 21: Delta-II Assembly Plan Build Delta-II On A CMM
Tune one quadrant.
Remove.
Tune the other quadrant.
Reinstall the first quadrant.
Center slides remain fixed.
Lay E flat on CMM.
Should give repeatable position to first quadrant after reinstallation.

## Slide 22: Delta-II Assembly Plan
Measure/Tune each half.
Measure assembled undulator.

## Slide 23: Fiducialize Strongbacks

## Slide 24: Adjust The Mechanical Axis
Move the blocks in order to move the magnet mechanical axis
to the fiducialized geometric axis of the strongback.

## Slide 25: Measure On Line Between Pointed Magnet Centers
Line radial taper less than 8 microns.
Line radial bow less than 18 microns.
Move blocks to tune quadrant.
‘Pointed magnets’ make the connection between CMM probe position and Hall probe position.
We purchased software to move the CMM in a straight line and generate triggers uniformly spaced.

## Slide 26: Assembled Undulator
Flats on E set Y and pitch.
Use tooling balls on E to set X and yaw.
Need tooling balls at ends to determine Y after assembly. (Green)
Planning new Hall probe for assembled Delta-II measurements.

## Slide 27: Hall Probe Measurements
Repeat the procedure used on the Delta undulator.

## Slide 28: Locate Magnetic Center
Center relative to probes.
Use linear field mode.
Measure Bx and By.
Can rotate probes to measure
at two points along x or y.
Assume field shape
is first term in field
expansion near the
center.
LCLS-TN-13-9

## Slide 29: Measure Probe Position
Transverse location from corner cube. Longitudinal location from interferometer.
Georg Gassner

## Slide 30
LCLS-II Undulators

## Slide 31: LCLS-II Undulator Tuning Overview
  22 SXR undulators (production plus spare)
  33 HXR undulators (production plus spare)
  3 measurement benches, 2 at SLAC, 1 at LBNL
  All final data sets will be done at SLAC with the undulator in the same orientation as in the tunnel
  For each undulator line, all final data sets will be done on the same bench and have the same set of systematic errors

## Slide 32: SXR Measurements With The HXU32 Prototype
Hall probe measurements are done with the Kugler bench.
Field integrals are measured with a long coil.

## Slide 33: HXR Measurements With The HGVPU Prototype
Hall probe measurements are done with the probe at the end of a large carbon fiber tube.
Long coil measurements will be done with a vertical coil.

## Slide 34: Hall Probe Speed Dependence
Signal frequency (Hz) Senis 118-114 peak field (T) Sentron 463 peak field (T)
0.33 1.276085 1.276314
0.67 1.274817 1.276309
1.33 1.273279 1.276304
2.67 1.269586 1.276329
Max rel. difference: 5·10-3 2·10-5
Isaac’s measurements
Yurii’s measurements
For now we will use our Sentron probe.
The new Senis probes seem to have a problem.

## Slide 35: K Fit Tests
HXU32
Halbach formula
Fit only odd points, find residuals for even and odd points

## Slide 36: K Fit Tests
HXU32
Extended Halbach formula
Fit only odd points, find residuals for even and odd points

## Slide 37: K Fit Tests
HXU32
Local polynomial fits
Fit only odd points, find residuals for even and odd points

## Slide 38: Phase Matching
Phase shims set the phase matching error to zero at one K value.
This lets us commission the undulators without phase shifters.
The entrance phase shifter makes the phase matching error in the undulator center a multiple of 2π.
The exit phase shifter makes the phase error at the exit cell boundary a multiple of 2π.
The actual phase shifter has the phase of the sum of the entrance plus previous exit phase shifters.
LCLS-TN-16-3

## Slide 39: Phase Matching Fits
We plan to use fits to determine the phase shifter settings between measured points.
HXU32 Measurements

## Slide 40: Magnet Radiation Damage Measurement Plan For The Tunnel
We are testing thin printed circuit coils that sit on the magnets with the poles sticking through holes in the kapton substrate.
The coil voltage is integrated to give the flux change as the gap is changed between two values. Changes in the flux measurement will indicate radiation damage.
S. Anderson
Kapton
1 m long
125 microns thick
2 layers
36 turns per pole
0.1 VS signal

## Slide 41: Repeatability Of Coil Measurements
S. Anderson

## Slide 42: Gap Encoder Radiation Damage
Set the undulator gap to the same value as a reference piece using capacitive sensors.
Record encoder readings.
If an encoder fails, repeat the procedure and set the new encoder offset to give the same encoder reading.

## Slide 43: Earth Field
Plan:
The undulators will be measured with the same orientation as they will have in the tunnel.
The fields in the tunnel will be mapped.
Correctors on the beam pipe will account for differences between the Earth field in the tunnel and in the lab.
We will construct a Helmholtz coil to simulate the Earth field differences between the tunnel and the lab.
We will determine current settings for the corrector coils based on the Helmholtz coil measurements and the tunnel map.
We will check the current settings in the tunnel using a portable vibrating wire system.

## Slide 44: Earth Field At SLAC And LBNL
SLAC
LBNL
LCLS-II
SLAC
Bx = +0.269 G
By = -0.343 G
LBNL
Bx = +0.126 G
By = -0.335 G
LBNL tuning won’t need major corrections at SLAC.
X = perpendicular to bench
Y = up
Z = along bench

## Slide 45: LCLS Tunnel Fields Measured For LCLS-I
0.1 G
-0.25 G
0.3 G
-0.6 G

## Slide 46: Correct Earth Field Difference With Beam Pipe Corrector Coil
I = 1 A
I = 3 A
Pole steel simulated by image currents.

## Slide 47: Beam Pipe And Magnet Array Heating
Temperature rise of “undulator jaws”
Temperature rise of beam pipe
Small thermistors on beam pipe and “undulator”
Need water cooled chamber so we don’t change the K value.
1 W/m corrector currents
plus
1 W/m heater (beam current)

## Slide 48: Helmholtz Coil To Simulate Earth Field Difference Between MMF And Tunnel
Coils
Truss structure
I1
I2
-I2
-I1
Two power supplies for two coils on diagonals.

## Slide 49: XFEL Earth Field Studies
Uwe Englisch

## Slide 50: Helmholtz Coil

## Slide 51: Field Integral Measurements In The Tunnel
Use beam pipe with corrector coils and a vibrating wire system at the beam position. The system is currently being tested.
Y. Levashov

## Slide 52: Phase Shifter Measurements
The SLAC measurements agree
with the Danfysik measurements.
All strength and field integral requirements were met.

## Slide 53: Phase Shifter Crosstalk Measurements
The edge of the phase shifter must be 14 cm from the last pole
of the undulator (using the 3 Gcm criterion, no mirror plates).
The current design meets this requirement.

## Slide 54: Summary
The Delta undulator has a gradient in K that must be accounted for when using multiple Deltas in a system.
Assembly tolerances on the Delta undulator are at the 8 micron level if we are to avoid large phase errors.
We have a plan for sharing the LCLS-II tuning work with LBNL.
We have started to fit the phase matching errors and the K value so that we can determine what density of measurements are required for the phase shifters and the undulators.
We are testing a scheme to monitor radiation damage in the tunnel.
We are dealing with the Earth’s magnetic field by measuring the undulators in the same orientation as in the tunnel, mapping differences in the Earth’s field between the MMF and the tunnel, and using corrector coils on the beam pipe to compensate the differences.