## Slide 1
Note: the LCLS Control System requires all dipole magnets powered in series to have positive IVB polynomials (i.e. BDES[kG-m] > 0 gives IDES[A] > 0). To that end, this dipole’s IVB (which is properly negative) has had its sign reversed both in Oracle and in EPICS.

## Slide 2
engineering type: 0.315D13.08-C
SLAC drawing number: Stangenes 731-30822-10 r2
SLAC magnet bar-code: 4596
vendor serial number: 30822-2
when measured: 18-Nov-2025 09:09:54
when reanalyzed: 03-Feb-2026
MAD assignment: BCX203282
Itrim = 0 (after bipolar standardize)
ΔB/B0 = -0.43% @ X = +9 mm < 0.5%
Imain = 13 A (wirevsx.ru9)
BCX203282: BL vs X
measured bipolar
operate unipolar
BTRM(BDES) = -0.07047 Amain
BTRM(IDES) = -0.28924 Atrim
for BL = 0 kG-m

## Slide 3
Harmonic tolerance: |ΔB/B0| < 0.5% @ |X| = 9 mm
wirevsx.ru9
wirevsx.r10
BCX203282 (Imain=13 A, Itrim=0)
BCX203282 (Imain=14.5 A, Itrim=0)

## Slide 4
Imain = 0 (after bipolar standardize)
wiredat.ru4
|Imain| < 10 A, Itrim = 0
(after bipolar standardize)
name: BCX203282T
# of main coil turns: 668
# of trim coil turns: 164
ratio of turns (Nmain/Ntrim): 4.0732
ratio of slopes (measured): 4.1488
IMMO (amps): [-6,6]
BMMO (main coil amps): [-1.4462,+1.4462]
fit to DOWN data
fit to UP data