## Slide 1
engineering type: 0.315D13.08-C
SLAC drawing number: Stangenes 731-30822-10 r2
SLAC magnet bar-code: 4598
vendor serial number: 30822-3
when measured: 03-Dec-2025 13:05:09
when reanalyzed: 03-Feb-2026
MAD assignment: BCX203286
Itrim = 0 (after bipolar standardize)
|ΔB/B0| = 0.38% @ X = +9 mm < 0.5%
Imain = 13 A (wirevsx.ru8)
BCX203286: BL vs X
measured bipolar
operate unipolar
BTRM(BDES) = -0.07089 Amain
BTRM(IDES) = -0.2939 Atrim
for BL = 0 kG-m

## Slide 2
Harmonic tolerance: |ΔB/B0| < 0.5% @ |X| ≤ 9 mm
BCX203286 (Imain=13 A, Itrim=0)
BCX203286 (Imain=14.5 A, Itrim=0)

## Slide 3
Imain = 0 (after bipolar standardize)
wiredat.ru4
|Imain| < 10 A, Itrim = 0
(after bipolar standardize)
name: BCX203286T
# of main coil turns: 668
# of trim coil turns: 164
ratio of turns (Nmain/Ntrim): 4.0732
ratio of slopes (measured): 4.1451
IMMO (amps): [-6,6]
BMMO (main coil amps): [-1.4474,+1.4474]
fit to DOWN data
fit to UP data

## Slide 4
bhvszdat.r11
BCX203286