Revision History

| Revision | Date Released | Description of Change |
| --- | --- | --- |
| R0 | January 6, 2022 | Original Release. |

This traveler is for the magnetic measurements of one type 1.085Q4.3, QE4, quadrupole magnet. These measurements are needed for the FACET-II Injector QE quadrupoles which are going to be run at +/-12 and +/- 30 amps, which is lower than their 200 A nominal operating currents.  The 1.085Q4.3 quad to be tested is a LCLS-II spare with barcode L204244.

Receiving:

The following information is to be noted upon receipt of the magnets by the SLAC MFD group:

| Received by (MMG initials): | SDA |
| --- | --- |
| Date received (dd-mmm-yyyy): | 1/15/2022 |
| SLAC barcode number: | L204244 |
| Vendor serial number on the magnet: | QE040 |

Preparation:

Magnet beam direction and polarity are already marked.

Fiducialization:

Fiducialization was done previously.

URL of on-line CMM fiducialization data (please modify or correct if necessary):

https://www-group.slac.stanford.edu/met/MagMeas/MAGDATA/LCLS-II/Fiducial%20Reports/L204244_Fiducial_Report.pdf

Magnetic Measurements:

Enter URL of on-line magnetic measurements data (please modify or correct if necessary):

http://www-group.slac.stanford.edu/met/MagMeas/MAGDATA/FACET_II/quad/1.085Q4.31_L204244

- The connection polarity of a “positive” quadrupole will maintained (see Figure 1):

Figure 1: A “QF” focuses the electron beam in the horizontal plane, represented as positive k-value in the MAD-deck.

- Connect the magnet terminals in the correct polarity as established above, to a bipolar power supply with maximum current I  30 A (intended power supply type in installation is MCOR 30).

- Train magnet in +/- 30 Range. Standardize the magnet, starting from -30 A then go through 15 full cycles from -30 A to +30 A, finally ending down at -30 A from which the first operating current will be reached, with a flat-top pause time (at both -30 and +30 A) of 10 seconds.  Use a “Linear” ramp rate of 3 A/sec, record the ramp rate, type and pause used.

| Ramp Type | Linear |
| --- | --- |
| Ramp rate used (A/sec): | 1 A/sec |
| Ramp Pause | 10 sec |

- After the standardize 3 cycles, measure the length-integrated field gradient, Gdl, at -30 A.  Trip magnet off at -30 amps and then standardize 3 cycles and measure the length-integrated field gradient, Gdl, at -30 again.

| Filename & run number of Gdl, at -30 A data: | Strdat.ru1, strplt.ru1 |
| --- | --- |

- Confirm the pole-tip field using a Hall probe at an excitation current of +30 A.  Run magnet to 0 and turn off supply.

| Hall probe pole-tip field at +30 A (mean of 4 poles): | 0.173 Tesla @ 30.00841 amps |
| --- | --- |

- Start full measurement run.  Run the magnet up to 30 A for 10 min for a thermal warmup, record the following data after the warmup.

| Ambient temperature (°C): | 16.4 °C |
| --- | --- |
| Final Coil temperature at 30 A maximum (°C): | 24.4 °C |
| Final Core temperature at 30 A maximum (°C): | 21.6 °C |
| Water Flow (GPM) | 1.6 gpm |
| ∆ Pressure (PSI) | 117 psi |

- Standardize the magnet, starting from -30 A then go through 3 full cycles from -30 A to +30 A Measure the length-integrated field gradient, Gdl, from -30 to +30 A in 3-A steps (21 ‘up’ measurements), and then back down from +30 A to -30 A in 3-A steps (20 ‘down’ measurements).

| Filename & run number of Gdl up & down data: | Strdat.ru2, strplt.ru2 |
| --- | --- |

- With rotating coil, measure the magnet harmonics at the +30 A current setting. Multipole values should be given as a percentage of the quadrupole moment evaluated at the probe radius.

| Filename & run number of harmonic data: | Hardat.ru2, harplt.ru2 |
| --- | --- |
| Probe radius used for harmonics (m): | 0.009347 m |
| Rotating Coil Designation (Name) | 0.75DQB26 |

- Connect the magnet terminals in the correct polarity as established above, to a bipolar power supply with maximum current I  12 A (intended power supply type in installation is MCOR 12).

- Train magnet in +/- 12 A Range. Standardize the magnet, starting from -12 A then go through 15 full cycles from -12A to +12 A, finally ending down at -12 A from which the first operating current will be reached, with a flat-top pause time (at both -12 and +12 A) of 10 seconds.  Use a “Linear” ramp rate of 1 A/sec. Record the ramp rate used, the ramp type and the pause duration.

| Ramp Type | Linear |
| --- | --- |
| Ramp rate used (A/sec): | 1 A/sec |
| Ramp Pause | 10 sec |

- After the standardize, measure the length-integrated field gradient, Gdl, at -12 and +12 A.

| Filename & run number of Gdl, at -12 and +12 A data: | Strdat.ru3, strplt.ru3 |
| --- | --- |

- Confirm the pole-tip field using a Hall probe at an excitation current of +12 A.  Run magnet to 0 and turn off supply.

| Hall probe pole-tip field at +12 A (mean of 4 poles): | 0.071 Tesla @ 12.00167 Amps |
| --- | --- |

- Standardize the magnet, starting from -12 A then go through 3 full cycles from -12 A to +12 A Measure the length-integrated field gradient, Gdl, from -12 to +12 A in 1-A steps (25 ‘up’ measurements), and then back down from +12 A to -12 A in 1-A steps (24 ‘down’ measurements).

| Filename & run number of Gdl up & down data: | Strdat.ru4, strplt.ru4 |
| --- | --- |

- With rotating coil, measure the magnet harmonics at the +12 A current setting. Multipole values should be given as a percentage of the quadrupole moment evaluated at the probe radius.

| Filename & run number of harmonic data: | Hardat.ru4, harplt.ru4 |
| --- | --- |
| Probe radius used for harmonics (m): | 0.009347 m |
| Rotating Coil Designation (Name) | 0.75DQB26 |

- Standardize the magnet, starting from 0 A then go through 10 full cycles from 0 A to +30 A Measure the length-integrated field gradient, Gdl, from 0 to +30 A in 1-A steps (31 ‘up’ measurements), and then back down from +30 A to 0 A in 1-A steps (30 ‘down’ measurements).

| Filename & run number of Gdl up & down data: | Strdat.ru5, strplt.ru5 |
| --- | --- |

- With rotating coil, measure the magnet harmonics at +12 and +30 A . Multipole values should be given as a percentage of the quadrupole moment evaluated at the probe radius.

| Filename & run number of harmonic data: | Hardat.ru5, harplt.ru5 |
| --- | --- |
| Probe radius used for harmonics (m): | 0.009347 m |
| Rotating Coil Designation (Name) | 0.75DQB26 |

- Measure the inductance and resistance of the magnet:

| Inductance of coil (mH): | 7.41 mH @ 100 Hz |
| --- | --- |
| Resistance of coil (Ohms): | 0.0834 Ohm |
| Coil temperature in degrees C | 23.1 oC |

- Upon completion of tests, email URL of on-line data to Mark Woodley. Mark Woodley will determine if the magnet data is accepted.  Upon acceptance of magnet data, analysis data will be placed in on-line data folder.

| Magnet data accepted and Analysis file(s) put into on-line data folder (initials): | SDA |
| --- | --- |

Enter URL of on-line magnetic measurements analysis data :

https://www-group.slac.stanford.edu/met/MagMeas/MAGDATA/FACET_II /QUAD/1.085Q4.31-L204244/1.085Q4.31_4244.pptx