## Test Procedure for the LBL BTA TQDB-B Q10 Quadrupole Magnets

(Revision 01)

This test procedure is intended to cover magnetic measurements for the LBL TQDB-B Q15 quadrupole magnets at SLAC.

Receiving:

The following information is to be noted upon beginning the measurements of the magnet by the SLAC MM group:

| Received by (initials): | SDA |
| --- | --- |
| Date put on Test stand (dd-mmm-yyyy): |  |
| Vendor Serial Number on the magnet: |  |
| Shipping damage (Y or N): |  |

Preparation:

Install magnet on test stand with rotating coil encoder on the upstream end of the magnet.

Procedures:

Ramp Rates and Wait Times:

Main Coil – 10 A/s, 20 seconds

Trim Coil – 2.5 A/s, 20 seconds

Linear Ramp

PROC01: Magnet Wash Cycle:

Ramp the magnet from 0 to 150 to 0 amps before starting measurements

Magnetic Measurements:

Enter URL of on-line magnetic measurements data:

http://www-group.slac.stanford.edu/met/MagMeas/MAGDATA/LBL/Quad/TQDB_B_0002_BTA_Q15

- Check that the quadrupole main and corrector connections are made so that a positive current produces fields as marked on the coils of the quadrupole for the quadrupole main , Bx trim and By trim.

- Confirm that the polarity for the main polarity markings on the magnet are correct:

| Quad Field polarity has been labeled correctly (initials): |  |
| --- | --- |

- Confirm that the polarity the Bx polarity markings on the magnet are correct:

| Bx trim lead connections are labeled correctly (initials): |  |
| --- | --- |

- Confirm that the polarity the By polarity markings on the magnet are correct:

| By lead connections are labeled correctly (initials): |  |
| --- | --- |

- Measure the inductance and resistance of the Quadruple coils:

| Inductance of Main coils at 100 Hz (mH): | 2.4334 mH |
| --- | --- |
| Resistance of Main coils (): | 0.1874 |
| Main Coil Temperature | 20.4 °C |

- Measure the inductance and resistance of the Bx trim coils:

| Inductance of Bx trim coils at 100 Hz (mH): | mH |
| --- | --- |
| Resistance of Bx trim coils (): |  |

- Measure the inductance and resistance of the By trim coils:

| Inductance of By trim coils at 100 Hz (mH): | mH |
| --- | --- |
| Resistance of By trim coils (): |  |

- Determine the number of current cycles at each current setpoint for main quadrupole harmonic integral reproducibility to < 3 units. Follow PROC01, then measure the length-integrated field gradient Gdl, at main current = 45 A for as many cycles as needed until main quadrupole harmonic integral reproducibility to < 3 units. Note the number of cycles needed.  This number of cycles will be used for each current setpoint.

| Rotating coil name | 1DQB26 |
| --- | --- |
| Rotating coil radius (m): | 0.012431 |
| Harmonics data file name: | Hardat.ru, Harplt.ru |
| Filename of Gdl data: | Strdat.ru, strplt.ru |
| # of Current Cycles need for < 3 unit reproducibility |  |

- Perform the wash cycle (PROC01), then measure the length-integrated field gradient and the field harmonics, for quadrupole current settings 45, 65, 92, 132 and 145 amps using the number of cycles in set 8.

| Harmonics data file name: | Hardat.ru, Harplt.ru |
| --- | --- |
| Filename of Gdl data: | Strdat.ru, strplt.ru |

- Set the main to the nominal calibrated current of 132 A, (∫Gdl  = 14.72 T).  Measure the harmonics with the Bx corrector trim at 0, +10, +25, -10, -25 and 0 amps.

| Harmonics data file name: | Hardat.ru, Harplt.ru |
| --- | --- |
| Filename of Gdl data: | Strdat.ru, strplt.ru |

- Set the main to the maximum calibrated current of 145 A, (∫Gdl  = 16.19 T).  Measure the harmonics with the Bx corrector trim at 0, +10, +25, -10, -25 and 0 amps.

| Harmonics data file name: | Hardat.ru, Harplt.ru |
| --- | --- |
| Filename of Gdl data: | Strdat.ru, strplt.ru |

- Set the main to the nominal calibrated current of 132 A, (∫Gdl  = 14.72 T).  Measure the harmonics with the By corrector trim at 0, +10, +25, -10, -25 and 0 amps.

| Harmonics data file name: | Hardat.ru, Harplt.ru |
| --- | --- |
| Filename of Gdl data: | Strdat.ru, strplt.ru |

- Set the main to the maximum calibrated current of 145 A, (∫Gdl  = 16.19 T).  Measure the harmonics with the By corrector trim at 0, +10, +25, -10, -25 and 0 amps.

| Harmonics data file name: | Hardat.ru, Harplt.ru |
| --- | --- |
| Filename of Gdl data: | Strdat.ru, strplt.ru |

- Upon completion of tests, send link to all data to Martin Johansson .  Acceptance of magnet should be given by email to Magnetic Measurements Group.

| Acceptance given by Martin Johansson, date |  |
| --- | --- |