## Slide 1: LCLS-II-HE SXU Short Prototype Measurements
Yurii Levashov, Engineer
LCLS-II-HE 2022 Director’s Status Review​
February 8-11, 2022

## Slide 2: Outline
Measurement set-up
Requirements
Tests:
  Evaluate magnetic structure design. Measure the prototype with calibrated Hall probe, analyze the fields. Measure field integrals vs. gap.
  Study shim signatures and saturation effects. Evaluate if shims are adequate for tuning.
  Study Hall probe measurements at Kugler bench and Hall probe calibration. Make sure the measurement accuracy meets the tolerances on undulator K (<20% of total tolerance).
  Study magnet crosstalk between SXU-HE undulator and new SXR-HE phase shifter. Find the minimum distance at which the change in field integrals is insignificant (< 3G-cm).
Summary
*Yurii Levashov, Zack Wolf “Magnetic Measurements of SXU-HE short prototype”, LCLS-TN-22-3.

## Slide 3: Measurement Set-up
Two 14 poles, 36cm long magnet structures.
1m long prototype frame, leftover from LCLS-II project.
No gap encoders.
SXR measurement bench.
Laser tracker for alignment.
Magnetic taper corrected with shims.
The same equipment as the LCLS-II project:
  Sentron 2-D Hall probe to measure fields.
  Stretched wire to measure field integrals (long coil, 150 turns).
Shims
Charge Question: 2.C, 3.B, 3.C

## Slide 4: Evaluation of magnetic structure design, requirements.
Requirement # | Parameter | Unit | Requirement | Verification | Test Results
PRD0049.4014 | Undulator period length | mm | 56 | Test | 56.031
PRD0049.4024 | Minimum operational magnetic gap | mm | 7.2 | Inspection | -
PRD0049.4025 | Maximum operational magnetic gap | mm | 33 | Inspection | -
PRD0049.4026 | On-Axis vertical effective field at min. oper. gap | T | >1.76 | Test | 1.85
PRD0049.4027 | Keff at minimum operational gap | - | >9.21 | Test/Analysis | 9.69
PRD0049.4033 | Undulator parameter tolerance ΔKeff/Keff | - | ±5.5×10-4 | Test/Analysis | -
PRD0049.4029 | Minimum operational K values | - | 1.51 | Test/ Analysis | 1.538 (at 33mm gap)
PRD0049.4034 | Horizontal K sextupole | 1/mm2 | <5×10-4 | Test/ Analysis | 3.5×10-4
PRD0049.4038 | Phase shake (rms) over Lcell | deg Xray | <5.0 | Test/ Analysis | 2.4 (Before tuning)
PRD0049.4040 | First field integral of By per cell (abs) | μTm | <50 | Test/ Analysis | 20 (After tuning)
PRD0049.4041 | Second field integral of By per cell (abs) | μTm2 | <200 | Test/ Analysis | -
PRD0049.4042 | First field integral of Bx per cell (abs) | μTm | <50 | Test/ Analysis | 20 (After tuning)
PRD0049.4043 | Second field integral of By per cell (abs) | μTm2 | <200 | Test/ Analysis | -
Charge Question: 1.A, 3.A, 3.B

## Slide 5: Measurements with Hall probe
  First measurement results (before tuning):
  Average Undulator Period = 0.056031 m ± 0.000070 m
  The K at 7.2mm gap = 9.69 ( 9.21 PRD); 5% stronger!
  Bx field is small, below 10G:
  No large magnetization errors in the magnets
  pole alignment is good
  magnet sorting is good
PRD0049.4014
PRD0049.4026
PRD0049.4027
Charge Question: 2.C, 3.B, 3.C

## Slide 6: Measurements with Hall probe; continued
  Phase shake = 2.4° (<5°)
  Magnetic field roll-off = 3.5·10-4 1/mm2 (< 5·10-4 )
  ~25µTm2 jumps after each end section in the same direction, 50µTm2 in total, agree with simulations done by LBNL (<200µTm2).
by D. Arbelaez
PRD0049.4034​
PRD0049.4038
Charge Question: 2.C, 3.B, 3.C

## Slide 7: Field integrals vs Gap
Only the first field integrals measured.
Background field integrals subtracted from all measurements.
Measurement error 3G-cm.
The sharp change in By field integral at ~ 9mm gap is due to pole saturation effects and magnetization errors.
Goal: Find shim signature with a similar shape.
Charge Question: 2.C, 3.B, 3.C

## Slide 8: Shim signatures; options for tuning
Two capacitive sensors
Move poles vertically
Small  3mm
x 3mm slugs
Large  8mm
x 3mm slugs
Add magnet material
Ferromagnetic shim
Steel shim 7.5 x 15 x 0.17mm glued on top of a magnet
Pole motion: The main dipole, the transverse dipole, normal quadrupole and phase errors.
Slugs = pole motion + Skew quadrupole.
Ferromagnetic shims = slugs + Sextupole components.
Combinations of different types of shims used to get the correct shim signature.
Charge Question: 2.C, 3.B, 3.C

## Slide 9: Main field signatures
Chart: Differences in shim signatures, By
| 7.2 | 7.5 | 9 | 12 | 15 | 20 | 25 | 30 | 40 | 50 | 75 | 100 | 125 | 150
H-slugs- small slugs | 17.5477 | 21.87 | 23 | 25.486 | 25.671 | 29.9027 | 33.5817 | 36.1533 | 35.1183 | 32.9373 | 36.129 | 38.0253 | 31.2177 | 21.5187 | 15.9293 | 11.3813
pole 50um- small slugs | 42.995 | 46.192 | 44.6173 | 36.9997 | 26.5843 | 12.8647 | 0.75933 | -11.1487 | -15.411 | -17.964 | -18.272 | -14.8357 | -8.0293 | -7.6587 | -4.6213 | -4.9947
steel shim - pole 50um | 6.5708 | 2.5813 | 11.5753 | 18.883 | 30.007 | 34.794 | 27.3797 | 17.711 | 12.625 | 7.0797 | 3.543 | 1.756 | -2.3743 | -4.8877 | -3.2 | 3.8223
Combination of pole motion with ferromagnetic shim produces a signature closest to the required field shape.
Chart: Shim signatures, By
| 7.2 | 7.5 | 8 | 9 | 10 | 12 | 15 | 20 | 25 | 30 | 40 | 50 | 75 | 100 | 125 | 150
4 small slugs | 88 | 81.7157 | 79.0707 | 76.261 | 68.7743 | 66.0403 | 55.9553 | 49.921 | 45.8117 | 42.364 | 30.471 | 22.8993 | 15.3087 | 11.8107 | 6.8213 | 7.4537
2 H-slugs | 105.5477 | 103.5857 | 101.747 | 95.943 | 89.537 | 86.0743 | 80.93 | 75.3013 | 66.6 | 60.9247 | 46.5263 | 33.3293 | 22.7507 | 18.835 | 96.12 | 96.12
pole 50um | 130.995 | 127.9077 | 123.688 | 113.2607 | 95.3587 | 78.905 | 56.7147 | 38.7723 | 30.4007 | 24.4 | 12.199 | 8.0637 | 7.2793 | 4.152 | 2.2 | 2.459
LCLS-II SXU | 96 | 110 | 88 | 72 | 52 | 37 | 25 | 20 | 10 | 6 | 2 | 1
Steel 7.5x15mm | 137.5658 | 130.489 | 135.2633 | 132.1437 | 125.3657 | 113.699 | 84.0943 | 56.4833 | 43.0257 | 31.4797 | 15.742 | 9.8197 | 4.905 | -0.73567 | -1 | 6.2813
No saturation effect visible at gaps < 8mm
Charge Question: 2.C, 3.B, 3.C

## Slide 10: Shim signatures for other field components
Chart: Bx field signatures
| 7.2 | 8 | 9 | 10 | 15 | 20 | 25 | 30 | 40 | 50 | 70 | 80 | 100 | 120 | 135 | 140 | 150
Pole cant | -67.4803 | -64.3863 | -61.8087 | -60.532 | -56 | -49.8007 | -40.566 | -32 | -24 | -19.1723 | -11.6627 | -9.5527 | -4.108 | -1 | 0.49167 | -0.015333
V-slugs | -7.148 | -4.853 | -1.882 | -3.3123 | -5 | -7 | -7.5677 | -12.4783 | -20.707 | -27.2937 | -34.626 | -38.5923 | -31.5553 | -23.9613 | -20.2797 | -12.1397
H-slugs | -18 | -6 | -5 | -4 | -5 | -6 | -3 | -9 | -13 | -19 | -11 | -8 | 1 | 10 | 15 | 16
Chart: Skew Quad signatures
| 7.2 | 8 | 9 | 10 | 15 | 20 | 25 | 30 | 40 | 50 | 70 | 80 | 100 | 120 | 135 | 140 | 150
4 small slugs | -6 | -9 | -5 | -18 | -5 | 3 | 1 | -4 | -2 | -2.5 | -4 | -6 | -7 | -5
2 H-slugs | -4 | -5 | -2.5 | -19 | -5 | 2.5 | -1 | -4.5 | 6 | 6 | 5 | 4 | 3 | 4 | 3 | 4 | 4
2 V-slugs | -4 | -8 | -15 | -22 | -9 | -6 | -9 | -11 | -9 | -8.5 | -8 | -8 | -5 | -1 | 2 | 4 | 7.5
Steel 7.5x15 | -79 | -62 | -60 | -56 | -30 | -6 | -2 | 3 | 3 | -1.5 | -2 | -3 | -3 | -2 | 0 | 1 | 3
Magnetic slugs are very weak.
Pole roll and ferromagnetic shims will be used.
Charge Question: 2.C, 3.B, 3.C

## Slide 11: Magnet/pole recess
Magnet
Pole
Steel shim 0.17mm thick + glue
Magnet recess 0.14mm
Needs to be ≥ 0.3mm
Poles moved by 0.150mm; checked on CMM.
The main field reduced by ~0.5% from 1.85T to 1.84T (4%).
No change in phase shake or field integrals.
Magnet/pole recess in the magnetic structure increased from 0.140mm to 0.300mm . Change request was generated.
Chart: Top Magnet/pole Y Alignment
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13
Magnets | -0.24502 | -0.25157 | -0.24577 | -0.30882 | -0.31912 | -0.31122 | -0.31392 | -0.31732 | -0.31187 | -0.31232 | -0.22932 | -0.23902 | -0.23472
Poles | -0.13882 | 0.023185 | 0.021735 | 0.014385 | 0.0086846 | 0.0055846 | 0.0071346 | 0.0049846 | 0.0066346 | 0.022785 | 0.0077346 | 0.0033346 | 0.020335 | -0.14652
Chart: Bottom Magnet/pole Y alignment
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13
Magnets | -0.20809 | -0.23674 | -0.25294 | -0.32369 | -0.32104 | -0.32184 | -0.32669 | -0.32294 | -0.32629 | -0.32119 | -0.28389 | -0.27254 | -0.20074
Poles | -0.14054 | 0.0069577 | 0.0062077 | 0.010208 | 0.028108 | 0.0076077 | 0.0074077 | 0.0097077 | 0.012008 | 0.0074077 | 0.011758 | 0.024558 | 0.012358 | -0.14429
CMM measurements
Charge Question: 1.A, 2.C, 3.B, 3.C

## Slide 12: Prototype tuning
The prototype field integrals were tuned to ±20µTm by applying combinations of shims.
PRD0049.4040
PRD0049.4042
by D. Arbelaez
by D. Arbelaez
Charge Question: 2.C, 3

## Slide 13: Tuning challenges
The pole adjustment works fine.
No access to the side screws near the frame vertical supports (~2 x 0.46m top & bottom).
A special tuning fixture is needed, like a fixture for HXUs.
No issues with installation of small slugs
Large slugs are very difficult to work with; slug diameter should be reduced.
Pole locking screw
Hex Tip
Knob
90° Gear
Charge Question: 2.C, 3.B, 3.C
LCLS-II SXU
Pole locking screws
Small slugs
HXU pole tuning fixture

## Slide 14: Hall probe calibration accuracy.
The Hall probe must be calibrated to higher field levels.
Two challenges:
  Probe output is very non-linear at fields above 1.5T. Cubic splines used to fit the calibration data.
  The calibration magnet saturates reducing the good field region. Probes to be positioned with high accuracy.
Project | Bund(T) | Bcalib(T)
LCLS-I | 1.25 | 1.5
LCLS-II | 1.62 | 1.7
LCLS-II-HE | 1.86 | 1.9
The calibration accuracy is ±5·10-5 T.
Charge Question: 2.C, 3.B, 3.C

## Slide 15: Hall probe measurement accuracy
The measurement accuracy checked by measuring the prototype repeatedly. The test consists of alternate Hall probe recalibrations and the prototype re-measurements.
The Hall probe measurement accuracy meets the LCLS-II-HE requirements (<1∙10-4).
Chart: Change in undulator K
| 44494.652777777781 | 44495.359722222223 | 44495.631944444445 | 44496.606944444444 | 44497.363888888889 | 44497.675000000003 | 44498.623611111114 | 44498.624305555553 | 44501.387499999997 | 44502.384722222225 | 44502.384722222225 | 44503.361111111109 | 44504.402777777781 | 44508.361111111109 | 44509.336805555555 | 44511
Undulator K | 0 | 1.1118 | 1.373 | 0.89055 | 1.0982 | 1.3174 | 2.82 | 2.8138 | 3.6105 | 3.0308 | 2.8085 | 3.057 | 2.1731 | 3.1964 | 2.8892 | 3.8778
Hall Probe calibrations
Gap of the prototype set to 7.3mm.
Ambient temperature variations (±0.1C°).
First 6 measurements during the gap initial drift ignored.
The deviations from the linear fit ±9∙10-5
PRD0049.4033*
Charge Question: 2.C, 3

## Slide 16: Crosstalk between undulator and the phase shifter
Chart: Change in Bx first field integral
| 25 | 20 | 15 | 12.5 | 10 | 10 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 13.5 | 12.5 | 11.5 | 12 | 10.5 | 9.5 | 10
 | 0.66667 | -1.3333 | -1.3333 | 1.6667 | -0.33333 | -1.3333 | 3.6667 | -0.33333 | 2.6667 | 1.6667 | 1.6667 | -1.3333 | -0.33333 | -0.33333 | 1.6667 | 1.6667 | 0.66667 | 1.6667 | 0.66667 | -0.33333 | 6.6667 | 0.66667 | 2.6667 | 0.66667
Rail
Slide
Distance Undulator/Phase shifter in tunnel = 14.4cm.
Insignificant cross-talk expected (< 2G-cm).
SXHEPS prototype
Chart: Change in By first field integral
| 25 | 20 | 15 | 10 | 10 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 13.5 | 12.5 | 11.5 | 12 | 10.5 | 10 | 10 | 10.5 | 11 | 11.5 | 12 | 12.5 | 13 | 13.5 | 14 | 14.5 | 15 | 16 | 17 | 18 | 19 | 20 | 25
I1Y | 1.1875 | 1.1875 | -1.8125 | -11.8125 | -7.8125 | -9.8125 | -3.8125 | -1.8125 | -3.8125 | 1.1875 | 1.1875 | 2.1875 | 3.1875 | -0.8125 | 1.1875 | 3.1875 | -0.8125 | -3.8125 | -3.8125 | -1.8125 | -6.8125 | -8.8125 | -14.8125 | -9.8125 | -6.8125 | -7.8125 | -5.8125 | -2.8125 | -2.8125 | -3.8125 | -1.8125 | -1.8125 | -2.8125 | -3.8125 | -1.8125 | -0.8125 | 0.1875 | -1.8125 | 0.1875
13cm
Undulator gap = 7.2mm
Phase shifter gap = 35mm
Charge Question: 2.C, 3

## Slide 17: Summary
The magnetic structure design meets the requirements. (The short prototype does not represent the full-scale device. Satisfaction is only partial).
Pole saturation is small.
Shim signatures measured. Field shimming options are sufficient for tuning the field integrals to tolerances.
Magnet/pole recess increased from 0.140mm to 0.300mm for ferromagnetic shim installation.
Hall probe measurements accuracy is <1∙10-4 , meets the requirements on K measurements.
Undulator/phase shifter cross-talk starts at 13cm. No cross-talk in the tunnel (14.4cm).
Next Pre-production prototype (Starts Q3 2022):
    Tune
    Measure full final dataset
    Temperature calibration
    Beam pipe corrector calibration