## Slide 1
BL @ Imain = 0 (after stdz): -23.10 G-m
engineering type: 0.315D14.39-C
SLAC drawing number: Everson Tesla Drawing # 53274
SLAC magnet bar-code: 4591
vendor serial number: 01
when measured: 24-Jul-2024 11:25:45
when analyzed: 09-Aug-2024
MAD assignment: BCXSS1

## Slide 2
quadrupole component removed
@ r= 10 mm
|b1/b0| = 0.057 [0.45] %
|b2/b0| = 0.053 [0.25] %
Imain = 75.02 A ; BL = 3.4 kG-m

## Slide 3
quadrupole component removed
@ r= 10 mm
|b1/b0| = 0.052 [0.45] %
|b2/b0| = 0.043 [0.25] %
Imain = 102.37 A ; BL = 4.5 kG-m

## Slide 4
Imain | b1/b0 % | b2/b0 % | b3/b0 %
-120 | 0.160 | 0.081 | 0.182
-60 | 0.131 | 0.075 | 0.162
-20 | 0.168 | 0.072 | 0.187
20 | 0.070 | 0.073 | 0.192
60 | 0.121 | 0.075 | 0.167
120 | 0.157 | 0.083 | 0.200
tolerance | 0.45 | 0.25 | -
Harmonic content vs Imain @ r = 1 cm
(rotating coil, rcoil = 3 mm)
hardat.r18

## Slide 5
after standardize
Imain = 0
slope @ Itrim = 0 : +10.3878 G-m/A
wiredat.ru9
after standardize
Imain = 102.7 A
slope @ Itrim = 0 : +7.0762 G-m/A
wiredat.r10
after degauss
Imain = 0
slope @ Itrim = 0 : +9.7072 G-m/A
wiredat.r15

## Slide 6
name: BCXSS1T
# of main coil turns: 80
# of trim coil turns: 18
ratio of turns (Nmain/Ntrim): 4.4444
ratio of slopes (measured): 4.6936
IMMO (amps): [-6,6]
BMMO (main coil amps): [-1.2783,+1.2783]
Imain = 0 (after degauss)
Itrim = 0 (|Imain| ≤ 95 A)
wiredat.ru2
wiredat.r15

## Slide 7
ramp type = 3_LINEAR
ramp rate = 15 A/s
settle time = 2 s
remnant BL = -5.56 +/- 0.76 μT-m
degauss time = 420 s
remnant BL taken from test plan

## Slide 8
bhvszdat.r16