## Slide 1
Note: the LCLS Control System requires all dipole magnets powered in series to have positive IVB polynomials (i.e. BDES[kG-m] > 0 gives IDES[A] > 0). To that end, this dipole’s IVB (which is properly negative) has had its sign reversed both in Oracle and in EPICS.

## Slide 2
engineering type: 0.315D14.39-C
SLAC drawing number: Everson Tesla Drawing # 53274
SLAC magnet bar-code: 4592
vendor serial number: 02
when measured: 05-May-2025 15:16:39
when analyzed: 10-May-2025
MAD assignment: BCXSS2
BL @ Imain = 0 (after stdz): 22.32 G-m

## Slide 3
quadrupole component removed
@ r= 10 mm
|b1/b0| = 0.020 [0.450] %
|b2/b0| = 0.044 [0.250] %
Imain = 75.02 A ; BL = -3.4 kG-m

## Slide 4
quadrupole component removed
@ r= 10 mm
|b1/b0| = 0.044 [0.450] %
|b2/b0| = 0.045 [0.250] %
Imain = 102.1 A ; BL = -4.5 kG-m

## Slide 5
after standardize
Imain = 0
slope @ Itrim = 0 : -9.7841 G-m/A
wiredat.ru9
after standardize
Imain = 102.2 A
slope @ Itrim = 0 : -6.8363 G-m/A
wiredat.r10
after degauss
Imain = 0
slope @ Itrim = 0 : -9.7295 G-m/A
wiredat.r13

## Slide 6
name: BCXSS2T
# of main coil turns: 80
# of trim coil turns: 18
ratio of turns (Nmain/Ntrim): 4.4444
ratio of slopes (measured): 4.6957
IMMO (amps): [-6,6]
BMMO (main coil amps): [-1.2777,+1.2777]
Imain = 0 (after degauss)
Itrim = 0 (|Imain| ≤ 95 A)
wiredat.ru2
wiredat.r13

## Slide 7
ramp type = 3_LINEAR
ramp rate = 15 A/s
settle time = 2 s
remnant BL = -23.12 +/- 0.80 μT-m
degauss time = 420 s
remnant BL: wiredat.r11