## Slide 1
engineering type: 0.315D14.39-C
SLAC drawing number: Everson Tesla Drawing # 53274
SLAC magnet bar-code: 4594
vendor serial number: 04
when measured: 21-May-2025 14:31:21
when analyzed: 30-May-2025
MAD assignment: BCXSS4
BL @ Imain = 0 (after stdz): -23.14 G-m

## Slide 2
quadrupole component removed
@ r= 10 mm
|b1/b0| = 0.001 [0.450] %
|b2/b0| = 0.041 [0.250] %
Imain = 75.03 A ; BL = 3.4 kG-m

## Slide 3
quadrupole component removed
@ r= 10 mm
|b1/b0| = 0.014 [0.450] %
|b2/b0| = 0.042 [0.250] %
Imain = 102.1 A ; BL = 4.5 kG-m

## Slide 4
after standardize
Imain = 0
slope @ Itrim = 0 : +9.8879 G-m/A
after standardize
Imain = 102.7 A
slope @ Itrim = 0 : +6.9992 G-m/A
after degauss
Imain = 0
slope @ Itrim = 0 : +9.6955 G-m/A
wiredat.r13
wiredat.ru9
wiredat.r10

## Slide 5
name: BCXSS4_Trim
# of main coil turns: 80
# of trim coil turns: 18
ratio of turns (Nmain/Ntrim): 4.4444
ratio of slopes (measured): 4.6958
IMMO (amps): [-6,6]
BMMO (main coil amps): [-1.2777,+1.2777]
Imain = 0 (after degauss)
Itrim = 0 (|Imain| ≤ 95 A)
wiredat.ru2
wiredat.r13

## Slide 6
ramp type = 3_LINEAR
ramp rate = 15 A/s
settle time = 2 s
remnant BL = -47.58 +/- 5.467 μT-m
degauss time = 420 s
remnant BL taken from test plan