## Slide 1: HE-SXU-000 Transportation & Storage test Results
Calibration measurements (October 12, 2022)
Transportation test (Test - October 20, 2022; Measurements - Nov.2, 2022)
Storage test (Dec. 6, 2022 – Jan. 17, 2023, Measurements - Jan. 20)
Temperature test (Jan. – Feb. 2023), in progress
Z. Wolf, Y. Levashov. LCLS-II-HE SXR Undulator Transportation Test Results. LCLS-TN-22-10

## Slide 2: HE-SXU-000 Tuning Results
  Summary:
  Field integrals tuned to tolerances in operational range 7.2mm to 33mm.
  Phase errors tuned well below tolerance in operational range.
  The K-value at 7.2mm gap is 9.7355 (~5% higher than required).
Chart: First field integrals
| 7.2 | 8 | 9 | 11 | 13 | 17 | 25 | 33 | 50 | 70 | 100
I1Y | -14.5383 | 2.521 | 28.0457 | 32.1403 | 30.1887 | 21.3947 | 9.0437 | 9.2033 | -9.8727 | -20.5697 | -28.895
Lower tolerance | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50 | -50
Higher tolerance | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50 | 50
I1X | -1.217 | -2.841 | 5.1377 | 18.2743 | 25.988 | 34.7273 | 28.5757 | 28.988 | -35.585 | -62.4003 | -66.8243
Chart: Second Field Integrals
| 7.2 | 8 | 9 | 11 | 13 | 17 | 25 | 33 | 50 | 70 | 100
Lower tolerance | -200 | -200 | -200 | -200 | -200 | -200 | -200 | -200 | -200 | -200 | -200
Higher tolerance | 200 | 200 | 200 | 200 | 200 | 200 | 200 | 200 | 200 | 200 | 200
I2X | 19.7013 | 21.2977 | 22.2223 | 30.1043 | 19.2517 | -0.81333 | -21.8877 | -12.8487 | -106.789 | -157.3453 | -173.598
I2Y | -58.327 | -41.0873 | -28.2127 | -23.1697 | -23.0197 | -15.828 | -18.6543 | -19.5647 | -34.323 | -52.0503 | -69.226

## Slide 3: Transportation Test
The undulator was moved to the tunnel and brought back the same day. After the transportation, the undulator was in the temperature stabilization area for 2 days, then moved to the Kugler bench for measurements.
Dataset 1 - Initial calibration
Dataset 2 – After the test

## Slide 4: Transportation Test, cont.
No Change!

## Slide 5: Storage test
The undulator was moved to the storage area in December 2022 and was back for measurements in 5 weeks.
Two temperature loggers were attached to the upper and lower jaws to record the temperature change.
Temp.
Loggers
The room heaters were running all the time by shipping/receiving .

## Slide 6: Storage test, continue
Dataset 2 - After transportation test
Dataset 3 – After the storage test

## Slide 7: Storage test, continue
No Change!