## Slide 1: Short prototype measurements (shim signatures)
Yurii Levashov
10/12/2021

## Slide 2: Outline
  Measurement set-up
  Field integrals’ gap dependence
  Options for shimming
  Main dipole field shims and corresponding signatures
  Skew quad shims
  Bx field shims
  Summary

## Slide 3: Measurement set-up
  Initial measurements by Sentron 2-D Hall probe.
  The K at 7.2mm gap = 9.69 ( 9.21 PRD); 5% stronger!
  Shim signature =
  difference in field integrals vs gap
  (Shim installed – No shim)
  Long coil used for measuring field integrals.

## Slide 4: SXU field integral gap dependence
Chart: Skew Quad
| 7.2 | 7.5 | 8 | 9 | 10 | 16 | 22
 | 207.6717 | 200.7567 | 178 | 141.5583 | 113.73 | 4 | 1
We expect a similar gap dependence of field integrals for SXU-HEs .
The source of the field errors to be investigated.
Goal: Find a shim signature with a close shape.
From LCLS-II experience the most challenging was the main field and the skew quad tuning.
SXU-017
SXU Short Prototype
Chart: Skew quad
| 7.2 | 7.5 | 8 | 9 | 10 | 15 | 20 | 30 | 40 | 50 | 100 | 150
 | 0.8 | 1 | -1.6 | -2.5 | -4 | -5.5 | -6.5 | -7 | -6 | -5 | -1.7 | -0.8

## Slide 5: Options for field integral tuning
Small  3mm
x 3mm slugs
Large  8mm
x 3mm slugs
Two capacitive sensors
Change gap by
moving poles
Add magnet material
Add iron shim
Steel shim 7.5 x 15 x 0.17mm glued on top of a magnet

## Slide 6: Main field shims
Large V-slugs
Pole shift
½ ∆G
Magnet
Pole
Steel shim
Small slugs
Pole
Large H-slugs
Magnet

## Slide 7: Main field signatures
Chart: Shim signatures, By Operational range
| 7.2 | 7.5 | 8 | 9 | 10 | 12 | 15 | 20 | 25 | 30 | 40 | 50 | 75 | 100 | 125 | 150
4 small slugs | 88 | 81.7157 | 79.0707 | 76.261 | 68.7743 | 66.0403 | 55.9553 | 49.921 | 45.8117 | 42.364 | 30.471 | 22.8993 | 15.3087 | 11.8107 | 6.8213 | 7.4537
2 H-slugs | 105.5477 | 103.5857 | 101.747 | 95.943 | 89.537 | 86.0743 | 80.93 | 75.3013 | 66.6 | 60.9247 | 46.5263 | 33.3293 | 22.7507 | 18.835 | 96.12 | 96.12
2 V slugs | 126.3357 | 118.9513 | 116.9 | 108.2747 | 97.9427 | 81.8613 | 68.9467 | 51.575 | 44.9957 | 34.534 | 19.8793 | 5.8117 | -6.0667 | -7.8627 | -6.8873
pole 50um | 130.995 | 127.9077 | 123.688 | 113.2607 | 95.3587 | 78.905 | 56.7147 | 38.7723 | 30.4007 | 24.4 | 12.199 | 8.0637 | 7.2793 | 4.152 | 2.2 | 2.459
LCLS-II SXU | 96 | 110 | 88 | 72 | 52 | 37 | 25 | 20 | 10 | 6 | 2 | 1
Steel 7.5x15mm | 137.5658 | 130.489 | 135.2633 | 132.1437 | 125.3657 | 113.699 | 84.0943 | 56.4833 | 43.0257 | 31.4797 | 15.742 | 9.8197 | 4.905 | -0.73567 | -1 | 6.2813
Chart: Differences in shim signatures, By
| 7.2 | 7.5 | 9 | 12 | 15 | 20 | 25 | 30 | 40 | 50 | 75 | 100 | 125 | 150
H-slugs- small slugs | 17.5477 | 21.87 | 23 | 25.486 | 25.671 | 29.9027 | 33.5817 | 36.1533 | 35.1183 | 32.9373 | 36.129 | 38.0253 | 31.2177 | 21.5187 | 15.9293 | 11.3813
V slugs - small slugs | 38.3357 | 37.2357 | 39 | 40.639 | 39.5003 | 31.9023 | 25.906 | 19.0257 | 5.7633 | 2.6317 | 4.063 | -3.02 | -9.497 | -17.8773 | -14.684 | -14.341
pole 50um- small slugs | 42.995 | 46.192 | 44.6173 | 36.9997 | 26.5843 | 12.8647 | 0.75933 | -11.1487 | -15.411 | -17.964 | -18.272 | -14.8357 | -8.0293 | -7.6587 | -4.6213 | -4.9947
steel shim - pole 50um | 6.5708 | 2.5813 | 11.5753 | 18.883 | 30.007 | 34.794 | 27.3797 | 17.711 | 12.625 | 7.0797 | 3.543 | 1.756 | -2.3743 | -4.8877 | -3.2 | 3.8223

## Slide 8: Skew quad shims
Chart: Skew Quad signatures
| 7.2 | 8 | 9 | 10 | 15 | 20 | 25 | 30 | 40 | 50 | 70 | 80 | 100 | 120 | 135 | 140 | 150
4 small slugs | -6 | -9 | -5 | -18 | -5 | 3 | 1 | -4 | -2 | -2.5 | -4 | -6 | -7 | -5
2 H-slugs | -4 | -5 | -2.5 | -19 | -5 | 2.5 | -1 | -4.5 | 6 | 6 | 5 | 4 | 3 | 4 | 3 | 4 | 4
2 V-slugs | -4 | -8 | -15 | -22 | -9 | -6 | -9 | -11 | -9 | -8.5 | -8 | -8 | -5 | -1 | 2 | 4 | 7.5
Steel 7.5x15 | -79 | -62 | -60 | -56 | -30 | -6 | -2 | 3 | 3 | -1.5 | -2 | -3 | -3 | -2 | 0 | 1 | 3
4 V-slugs | -10 | -12 | -27 | -42 | -33 | -25 | -24 | -23 | -23 | -22 | -17 | -13 | -8 | -6 | -3 | -1 | 1
Large V-slugs
Small slugs
Pole
Large H-slugs
Magnet
Magnet
Pole
Steel shim

## Slide 9: Bx field shims
Chart: Bx field signatures
| 7.2 | 8 | 9 | 10 | 15 | 20 | 25 | 30 | 40 | 50 | 70 | 80 | 100 | 120 | 135 | 140 | 150
Pole cant | -67.4803 | -64.3863 | -61.8087 | -60.532 | -56 | -49.8007 | -40.566 | -32 | -24 | -19.1723 | -11.6627 | -9.5527 | -4.108 | -1 | 0.49167 | -0.015333
V-slugs | -7.148 | -4.853 | -1.882 | -3.3123 | -5 | -7 | -7.5677 | -12.4783 | -20.707 | -27.2937 | -34.626 | -38.5923 | -31.5553 | -23.9613 | -20.2797 | -12.1397
H-slugs | -18 | -6 | -5 | -4 | -5 | -6 | -3 | -9 | -13 | -19 | -11 | -8 | 1 | 10 | 15 | 16
Pole roll
Pole
Large H-slugs
Magnet
Large V-slugs

## Slide 10: Summary
The main field on axis at 7.2mm gap is stronger than needed by ~5%.
The field integral gap dependence is similar to SXU undulators.
The pole saturation is small.
Skew quadrupole shims are weak. Steel shims are desirable for having correct shim signatures for the main field and the skew quadrupole.
To have steel shim option for tuning, the magnet should be recessed w.r.t. the poles by 0.250mm, minimum.
To test if the bigger magnet recess does not reduce the main field too much and if it will reduce the effect of magnetization errors , the following is planned:
Remove magnet arrays from the strongbacks (1 day).
Move all poles up by 150µm (technician training in pole adjustments, 1 day).
Check pole/magnet alignment on CMM (CMM programming for SXU-HE, 3 days).
Re-install magnet arrays/re-measure with Hall probe (~1 week).