## Slide 1: LCLS-II-HE SXU short prototype measurements
Yurii Levashov 2/11/2022

## Slide 2: Motivation & Scope
New, 56 mm period magnet structure were designed by LBNL. The structures were optimized for max magnetic field and minimum pole saturation. The goal was to evaluate the new magnetic design, confirm that it meets all the PRD tolerances on magnetic field.
Get hands-on experience working with the new magnetic structure; be prepared in advance for the full-scale prototype tuning.
Tests:
  Evaluate magnetic structure design. Measure the prototype with calibrated Hall probe, analyze the fields. Measure field integrals vs. gap.
  Study shim signatures and saturation effects. Evaluate if shims are adequate for tuning.
  Study Hall probe measurements at Kugler bench and Hall probe calibration. Make sure the measurement accuracy meets the tolerances on undulator K (<20% of total tolerance).
  Study magnet crosstalk between SXU undulator and new SXR phase shifter. Find the minimum distance at which the change in field integrals is insignificant.

## Slide 3: Set-up
Two 14 poles, 36cm long magnet structures were ordered, assembled and mechanically tested at LBNL.
At SLAC we had the structures assembled on a 1m long prototype frame, leftover from LSLC-II project, to make a short prototype of LCLS-HE undulator.
No gap encoders; the motor encoders’ offsets set at 7.2mm gap with the use of a ceramic block. Accuracy of setting the gap with motor encoders enough for the tests.
The prototype tested at SLAC MMF on SXU measurement bench. Magnet structures aligned to each other and to the bench granite with the use of a laser tracker.
The taper in the magnetic field along the beam axis corrected by inserting shims in between the magnets’ keepers and supporting strongbacks.
The same Sentron 2-D Hall probe used to measure fields, as for LCLS-II SXUs.
Stretched wire technique used to measure field integrals. The same device used for LCLS-I and LCLS-II field integral measurements (long coil, 150 turns).

## Slide 4: Evaluation of magnetic structure design
  Initial measurements (before tuning):
  The K at 7.2mm gap = 9.69 ( 9.21 PRD); 5% stronger!
  Bx field is small, below 10G. No large magnetization errors in the magnets, the pole alignment is good, magnet sorting is good.
  Phase shake = 2.4° (<5°)
  Magnetic field roll-off = 3.5·10-4 1/mm2 (< 5·10-4 )
  Saturation starts below 8mm gap and is small (~ 2·10-3).
~25µTm2 jumps after each end section in the same direction, ~50µTm2 in total, agree with simulations done by LBNL (<200µTm2).

## Slide 5: Field integrals vs Gap
The device is short, only the first field integrals measured.
Background field integrals measured without prototype and subtracted from all measurements.
Measurement errors estimated to be below 5G-cm.
The sharp change in By field integral at ~ 9mm gap is due to pole saturation effects and magnetization errors.
Goal: Find shim signature with a similar shape.

## Slide 6: Shim signatures; options for tuning
Small  3mm
x 3mm slugs
Large  8mm
x 3mm slugs
Two capacitive sensors
Move poles vertically
Add magnet material
Ferromagnetic shim
Steel shim 7.5 x 15 x 0.17mm glued on top of a magnet
The main dipole, the transverse dipole and normal quadrupole could be corrected by pole shifts. With slugs, skew quadrupole fields could be corrected. Sextupole components could be corrected by ferromagnetic shims only.
Combinations of different types of shims used to get the correct shim signature.
Shim signature shows how installation of a shim changes field integrals vs gap. (How a shim strength changes vs. gap.) Ideal shim signature has the same gap dependence as field integral to be tuned.

## Slide 7: Main field signatures
Chart: Shim signatures, By Operational range
| 7.2 | 7.5 | 8 | 9 | 10 | 12 | 15 | 20 | 25 | 30 | 40 | 50 | 75 | 100 | 125 | 150
4 small slugs | 88 | 81.7157 | 79.0707 | 76.261 | 68.7743 | 66.0403 | 55.9553 | 49.921 | 45.8117 | 42.364 | 30.471 | 22.8993 | 15.3087 | 11.8107 | 6.8213 | 7.4537
2 H-slugs | 105.5477 | 103.5857 | 101.747 | 95.943 | 89.537 | 86.0743 | 80.93 | 75.3013 | 66.6 | 60.9247 | 46.5263 | 33.3293 | 22.7507 | 18.835 | 96.12 | 96.12
2 V slugs | 126.3357 | 118.9513 | 116.9 | 108.2747 | 97.9427 | 81.8613 | 68.9467 | 51.575 | 44.9957 | 34.534 | 19.8793 | 5.8117 | -6.0667 | -7.8627 | -6.8873
pole 50um | 130.995 | 127.9077 | 123.688 | 113.2607 | 95.3587 | 78.905 | 56.7147 | 38.7723 | 30.4007 | 24.4 | 12.199 | 8.0637 | 7.2793 | 4.152 | 2.2 | 2.459
Steel 7.5x15mm | 137.5658 | 130.489 | 135.2633 | 132.1437 | 125.3657 | 113.699 | 84.0943 | 56.4833 | 43.0257 | 31.4797 | 15.742 | 9.8197 | 4.905 | -0.73567 | -1 | 6.2813
Chart: Differences in shim signatures, By
| 7.2 | 7.5 | 9 | 12 | 15 | 20 | 25 | 30 | 40 | 50 | 75 | 100 | 125 | 150
H-slugs- small slugs | 17.5477 | 21.87 | 23 | 25.486 | 25.671 | 29.9027 | 33.5817 | 36.1533 | 35.1183 | 32.9373 | 36.129 | 38.0253 | 31.2177 | 21.5187 | 15.9293 | 11.3813
pole 50um- small slugs | 42.995 | 46.192 | 44.6173 | 36.9997 | 26.5843 | 12.8647 | 0.75933 | -11.1487 | -15.411 | -17.964 | -18.272 | -14.8357 | -8.0293 | -7.6587 | -4.6213 | -4.9947
steel shim - pole 50um | 6.5708 | 2.5813 | 11.5753 | 18.883 | 30.007 | 34.794 | 27.3797 | 17.711 | 12.625 | 7.0797 | 3.543 | 1.756 | -2.3743 | -4.8877 | -3.2 | 3.8223
Combination of pole motion with ferromagnetic shim produces a signature closest to the required field shape.
Magnet recess in the magnetic structure increased from 0.140mm to 0.300mm, so that the installation of the ferromagnetic shims will not reduce the gap.
The main field reduced by ~0.5% from 1.85T to 1.84T (4%).
No change in phase shake or field integrals.

## Slide 8: Shim signatures for other field components.
Chart: Bx field signatures
| 7.2 | 8 | 9 | 10 | 15 | 20 | 25 | 30 | 40 | 50 | 70 | 80 | 100 | 120 | 135 | 140 | 150
Pole cant | -67.4803 | -64.3863 | -61.8087 | -60.532 | -56 | -49.8007 | -40.566 | -32 | -24 | -19.1723 | -11.6627 | -9.5527 | -4.108 | -1 | 0.49167 | -0.015333
V-slugs | -7.148 | -4.853 | -1.882 | -3.3123 | -5 | -7 | -7.5677 | -12.4783 | -20.707 | -27.2937 | -34.626 | -38.5923 | -31.5553 | -23.9613 | -20.2797 | -12.1397
H-slugs | -18 | -6 | -5 | -4 | -5 | -6 | -3 | -9 | -13 | -19 | -11 | -8 | 1 | 10 | 15 | 16
Chart: Skew Quad signatures
| 7.2 | 8 | 9 | 10 | 15 | 20 | 25 | 30 | 40 | 50 | 70 | 80 | 100 | 120 | 135 | 140 | 150
4 small slugs | -6 | -9 | -5 | -18 | -5 | 3 | 1 | -4 | -2 | -2.5 | -4 | -6 | -7 | -5
2 H-slugs | -4 | -5 | -2.5 | -19 | -5 | 2.5 | -1 | -4.5 | 6 | 6 | 5 | 4 | 3 | 4 | 3 | 4 | 4
2 V-slugs | -4 | -8 | -15 | -22 | -9 | -6 | -9 | -11 | -9 | -8.5 | -8 | -8 | -5 | -1 | 2 | 4 | 7.5
Steel 7.5x15 | -79 | -62 | -60 | -56 | -30 | -6 | -2 | 3 | 3 | -1.5 | -2 | -3 | -3 | -2 | 0 | 1 | 3
The prototype field integrals were tuned to tolerances (±10µTm) by applying combinations of shims.

## Slide 9: Hall probe calibration accuracy.
The Hall probe has to be calibrated to higher field levels.
Chart: Calibration Magnet Field roll-off 35mm Gap
| -20 | -15 | -10 | -5 | 0 | 5 | 10 | 15 | 20
1.9T | -2.9001 | -1.7348 | -0.81202 | -0.23201 | 0 | -0.094911 | -0.52729 | -1.2022 | -2.2462
1.7T | -1.5863 | -0.93107 | -0.39657 | -0.13219 | 0 | -0.045979 | -0.27587 | -0.62646 | -1.2127
1.5T | -0.20992 | -0.12866 | -0.088032 | -0.067717 | -0.054174 | -0.04063 | -0.013543 | 0 | 0
1T | 0.17 | 0.13 | 0.06 | 0.01 | 0 | 0.02 | 0.08 | 0.16 | 0.26
The calibration magnet saturates reducing the good field region. Probes to be positioned with high accuracy.
Project | Bund(T) | Bcalib(T)
LCLS-I | 1.25 | 1.5
LCLS-II | 1.62 | 1.7
LCLS-II-HE | 1.76 | 1.9
Probe output is very non-linear at fields above 1.5T. Cubic splines used to fit the calibration data.
NMR Probe
Hall Probe
No room to reduce the gap
The calibration accuracy is ±5·10-5 T.

## Slide 10: Hall probe measurement accuracy
The measurement accuracy checked by measuring the prototype repeatedly. The test consists of alternate Hall probe recalibrations and the prototype re-measurements.
The gap of the prototype set to 7.3mm and was not changed during the test. The change in the gap was due to the gap drift under the magnetic field force and the gap change due to ambient temperature variations.
Chart: Change in undulator K
| 44494.652777777781 | 44495.359722222223 | 44495.631944444445 | 44496.606944444444 | 44497.363888888889 | 44497.675000000003 | 44498.623611111114 | 44498.624305555553 | 44501.387499999997 | 44502.384722222225 | 44502.384722222225 | 44503.361111111109 | 44504.402777777781 | 44508.361111111109 | 44509.336805555555 | 44511
Undulator K | 0 | 1.1118 | 1.373 | 0.89055 | 1.0982 | 1.3174 | 2.82 | 2.8138 | 3.6105 | 3.0308 | 2.8085 | 3.057 | 2.1731 | 3.1964 | 2.8892 | 3.8778
First few measurements during the gap initial drift ignored. The deviations from the linear fit are less than ±9∙10-5 .
The measurement errors include the calibration, probe alignment, and the temperature gap drift errors.
The Hall probe measurement accuracy meets the LCLS-HE requirements (<1∙10-4).
Calibrations

## Slide 11: Crosstalk between undulator and the phase shifter.
If the distance between an undulator and a phase shifter is too small the field integrals will change with gap change.
Set the prototype and a phase shifter next to each other at different spacing between them and measure field integrals vs. gap to find the minimum distance at which the change in field integrals is insignificant.

## Slide 12: Summary.
The magnetic structure design meets the requirements. The short prototype does not represent the full scale device. Only partial satisfaction of these requirements could be stated.
Field shimming options are sufficient for tuning the field integrals.
Magnet/pole recess increased from 0.140mm to 0.300mm for ferromagnetic shims installation.
Hall probe measurements accuracy is <1∙10-4 , meets the requirements on K measurements.
Undulator/phase shifter cross-talk starts at cm.