## Slide 1
Chart: CMM Measurements
Corrected by shimming
Touch Probe after pitch correction
Frame support postions
CMMM measurements agree well with touch probe measurements

## Slide 2
Touch Probe after ends correction
Touch Probe final
No Access to screws to correct
Corrected to remove gap waviness

## Slide 3
Chart: Center Line Shift
| 7.5 | 9 | 11 | 13 | 17 | 25 | 33
Touch Probe | -35 | -6 | 4 | 18 | 48 | 58 | 76
Chart: 
| 7.5 | 9 | 11 | 13 | 17 | 25 | 33
Pitch | 2 | 0 | -1 | -4 | -5 | -6 | -6.5
Chart: Change in average gap
| 7.5 | 9 | 11 | 13 | 17 | 25 | 33
Gap Change | 0 | 0.006 | 0.01 | 0.013 | 0.014 | 0.014 | 0.014
Chart: Taper
| 7.5 | 9 | 11 | 13 | 17 | 25 | 33
Taper | 2 | -1 | 0.5 | -3 | -1 | -2.5 | -0.9
1. Gap center shift is
~ 100µm as expected.
2. Pitch is small.
3. Change in average gap indicates the strongbacks bending.
4. Taper is at the measurement error level.