## Slide 1
First measurements
of HE-SXU-023
Yurii Levashov
06/04/2024

## Slide 2: Outline
Mechanical measurement and alignment
Midplane encoder’s check
Magnetic measurements

## Slide 3: Mechanical measurements with touch probe, Kugler bench.
After initial alignment on Kugler bench, first measurement
Corrected by motor encoders’ offsets

## Slide 4: Mechanical measurements with touch probe, continue.
Pitch and taper corrected by changing encoders’ offsets.
Removing shims
between magnet
keepers and strongback
~25µm or less

## Slide 5: Mechanical measurements with touch probe, continue.
Final measurement
Jaws straight to
± 12µm.
Gap deviations
± 20µm.

## Slide 6: Midplane encoder’s test
Midplane encoders used to correct for undulator midplane shift vs. gap. By changing the gap offset.
Gap offset is calculated based on the difference between midplane encoders and gap encoders. Offset is corrected by a factor of 1.40
Midplane encoders
LCLS-TN-22-8 Z. Wolf, "HE-SXU Midplane Position Errors and Corrections." (October 2022)

## Slide 7: Midplane shift vs. gap
Chart: MA Vertical shift vs.Gap
| 7.5 | 9 | 10 | 11 | 13 | 17 | 25 | 33
Touch probe (no corrections) | -52 | -28 | 0 | -20 | 11 | 40 | 52 | 70
HP dense (No correction) | -79 | -25 | -12 | -6 | 0 | -2 | -7 | -12 | 5 | 11 | 19 | 48 | 55 | 73
Chart: MA Vertical shift vs.Gap
| 7.5 | 9 | 10 | 11 | 13 | 17 | 25 | 33
Touch probe (no corrections) | -52 | -28 | 0 | -20 | 11 | 40 | 52 | 70
HP dense (No correction) | -79 | -25 | -12 | -6 | 0 | -2 | -7 | -12 | 5 | 11 | 19 | 48 | 55 | 73
LT Roll | -0.175 | -0.108 | -0.045 | 0 | 0.035 | 0.061 | 0.098 | 0.148 | 0.199 | 0.216
Chart: MA Vertical shift vs.Gap
| 7.2 | 7.5 | 9 | 10 | 11 | 13 | 17 | 25 | 33
Hall Probe (corrected) | -39 | -23 | -7 | 0 | 0 | 0 | 1 | 4 | 0
Hall Probe repeat(Corrected) | -23 | -7 | 0 | 0 | 2 | 5 | 2
HP repeat (correction 1.4) | -46 | -4 | -2 | 4 | 2 | 0 | -1 | 1 | -3
Chart: MA Vertical shift vs.Gap
| 7.2 | 7.5 | 9 | 10 | 11 | 13 | 17 | 25 | 33
Hall Probe (corrected) | -39 | -23 | -7 | 0 | 0 | 0 | 1 | 4 | 0
Hall Probe repeat(Corrected) | -23 | -7 | 0 | 0 | 2 | 5 | 2
HP after Cycling | -7 | -4 | -1 | 0 | 6 | 3 | 0 | 0 | 2 | -2
HP repeat (correction 1.4) | -46 | -4 | -2 | 4 | 2 | 0 | -1 | 1 | -3
Chart: MA Vertical shift vs.Gap
| 7.2 | 7.5 | 9 | 10 | 11 | 13 | 17 | 25 | 33
Hall Probe (corrected) | -39 | -23 | -7 | 0 | 0 | 0 | 1 | 4 | 0
Hall Probe repeat(Corrected) | -23 | -7 | 0 | 0 | 2 | 5 | 2
HP repeat (correction 1.4) | -46 | -4 | -2 | 4 | 2 | 0 | -1 | 1 | -3
HP after Cycling | -7 | -4 | -1 | 0 | 6 | 3 | 0 | 0 | 2 | -2
HP after cycling, repeat | -32 | -5 | -3 | -2 | 5 | 3 | 0 | 0 | 1 | -3
Chart: HE-SXU-000 Center Line Shift (Corrected)
| 7.2 | 9 | 11 | 13 | 17 | 25 | 33
Offset Corrected | -6 | -10 | -2 | 2 | 10 | 14 | 16
Not real, due to bad data fit.
Backlash in the drive system when motion direction reverses.

## Slide 8: Magnetic Measurements, Field integrals
Bx
By
First integral tolerance:
±50 µTm
Second integral tolerance:
±200 µTm2

## Slide 9: Magnetic Measurements, Hall probe, 10mm gap
RMS 0.12%,
tolerance <0.25%
RMS 2.8°,
tolerance < 5°
Excellent sorting and mechanical assembly!

## Slide 10: Tuning

## Slide 11: Tuning, continue
First round |  |  |  |  | 
6/3/2024 |  |  |  |  | 
Location | Type of correction | Lower-Bench(Blue) | Lower-Aisle(Red) | Upper-Bench(Blue) | Upper-Aisle(Red)
Pole :5- Module#1 :Pole :5 | Pole Adjust: TB | +6 um | -6 um | -6 um | +6 um
Pole :6- Module#1 :Pole :6 | Pole Adjust: MB | +25 um | +25 um | +25 um | +25 um
Pole :18- Module#1 :Pole :18 | Pole canting: NQ | -25 um | +25 um | -25 um | +25 um
Pole :21- Module#1 :Pole :21 | Pole Adjust: MB | -25 um | -25 um | -25 um | -25 um
Pole :22- Module#1 :Pole :22 | Pole Adjust: MB | -25 um | -25 um | -25 um | -25 um
Pole :23- Module#1 :Pole :23 | Pole canting: NQ | +25 um | -25 um | +25 um | -25 um
Pole :45- Module#2 :Pole :3 | Pole Adjust: MB | +25 um | +25 um | +25 um | +25 um
Pole :46- Module#2 :Pole :4 | Pole Adjust: MB | +25 um | +25 um | +25 um | +25 um
Pole :80- Module#2 :Pole :38 | Pole canting: NQ | +25 um | -25 um | +25 um | -25 um
Pole :82- Module#3 :Pole :2 | Pole Adjust: MB | +25 um | +25 um | +25 um | +25 um
Pole :83- Module#3 :Pole :3 | Pole Adjust: MB | +25 um | +25 um | +25 um | +25 um
Pole :84- Module#3 :Pole :4 | Pole canting: NQ | -25 um | +25 um | -25 um | +25 um
Pole :86- Module#3 :Pole :6 | Pole Adjust: MB | +25 um | +25 um | +25 um | +25 um
Pole :87- Module#3 :Pole :7 | Pole Adjust: MB | +25 um | +25 um | +25 um | +25 um
Pole :91- Module#3 :Pole :11 | Pole Adjust: MB | +19 um | +19 um | +19 um | +19 um
Pole :94- Module#3 :Pole :14 | Pole canting: NQ | -25 um | +25 um | -25 um | +25 um
First round |  |  |  |  | 
6/3/2024 |  |  |  |  | 
Location | Type of correction | Lower-Bench(Blue) | Lower-Aisle(Red) | Upper-Bench(Blue) | Upper-Aisle(Red)
Pole :7- Module#1 :Pole :7 | Small slugs: MB | 3 slugs, away | 3 slugs, away | 3 slugs, towards | 3 slugs, towards
Pole :15- Module#1 :Pole :15 | Slugs: TB Ver | 2 towards | 2 away | 2 towards | 2 away
Pole :46- Module#2 :Pole :4 | Small slugs: MB | 3 slugs, away | 3 slugs, away | 3 slugs, towards | 3 slugs, towards
Pole :73- Module#2 :Pole :31 | Slugs: TB Ver | 2 towards | 2 away | 2 towards | 2 away
Pole :98- Module#3 :Pole :18 | Slugs: TB Ver | 2 towards | 2 away | 2 towards | 2 away