SLAC Magnetic Measurement Plan and Traveler for FACET-II Injector Laser Heater Undulator, (Nov 7, 2022)

This traveler covers magnetic measurements of the laser-heater permanent magnet planar undulator needed for FACET-II.  The undulator magnet is about 0.5 m long with 9 periods, each with a 5.4-cm length.

Receiving:

The following information is to be noted upon receipt of the magnets by the SLAC MFD group:

| Received by (initials): | YL |
| --- | --- |
| Date received (dd-mmm-yyyy): | 11/11/2022 |
| Serial number on the magnet (if available): | LHU-0003 |

Preparation:

A beam direction arrow, with text “beam direction”, is to be applied to the top and/or connector side of the magnet with a sticker supplied by FACET-II (Doug McCormick will determine the direction).

| Beam-direction arrow in place (initials): | YL |
| --- | --- |

Magnetic Measurements:

Enter URL of on-line magnetic measurements data (please modify or correct if necessary):

V:\MET\MagServe\MagData\LCLS-II-HE\Undulator\LHU-003\DATASET0001\Tuning\Z Scans\Run …

- Allow the magnet to come to an equilibrium temperature.

- At a gap setting corresponding to undulator K of 1.17, determine the magnetic undulator axis and fiducialize the magnet to external tooling balls so that the undulator can be aligned to the beam coordinates along its central axis to an accuracy of <0.5 mm horizontally and vertically. Use ceramic gage blocks to set the gap at magnet #5. The gap at K = 1.17 should be about 39 mm (TBD).

- After setting the undulator to a K of 1.17 measure and tune the undulator first field integrals to ±0.096 G-m (9.6 µTm). Measure and tune the undulator second field integrals to ±0.20 G-m2 (20µTm2).

- Measure the undulator K value and the precise gap height settings associated with the following K values ( 1.0, 1.1, 1.17, 1.3, and 1.4).

| Configuration | Measured K Km | Measured gap (mm) | Run # |
| --- | --- | --- | --- |
| K 1.0 | 0.98469 | 40 | 31 |
| K 1.1 | 1.11677 | 38 | 30 |
| K 1.17 | 1.18793 | 37 | 37 |
| K 1.3 | 1.29953 | 35.5 | 38 |
| K 1.4 | 1.41642 | 34 | 29 |
| K=1.17 | 1.17078 | 37.245 | 42 |

Measurement temperature Tm= 20.065C°. The K in the tunnel Kt = Km(1-10-3(T-Tm)).

- Measure the 1st and 2nd field integrals all along the undulator length for both the horizontal and vertical plane in the following conditions in the mid-plane. Directory name:

V:\MET\MagServe\MagData\LCLS-II-HE\Undulator\LHU-003\DATASET0001\Tuning\Stretched wire\Run 26

| Full gap height set for K = 1.17 (x = 15 mm): | Filename: integrals_summary.txt |
| --- | --- |
| Full gap height set for K = 1.17 (x = 10 mm): | Filename: integrals_summary.txt |
| Full gap height set for K = 1.17 (x = 5 mm): | Filename: integrals_summary.txt |
| Full gap height set for K = 1.17 (x = 0): | Filename: integrals_summary.txt |
| Full gap height set for K = 1.17 (x = 5 mm): | Filename: integrals_summary.txt |
| Full gap height set for K = 1.17(x = 10 mm): | Filename: integrals_summary.txt |
| Full gap height set for K = 1.17 (x = 15 mm): | Filename: integrals_summary.txt |

- Upon completion of tests, email URL of on-line data to Carsten Hast. Carsten Hast will determine if the magnet is accepted.  Upon acceptance of magnet, analysis data will be placed in on-line data folder.

| Magnet accepted and Analysis file(s) put into on-line data folder (initials): |  |
| --- | --- |
| Assigned beamline location (MAD-deck name): |  |