## LCLS-II Traveler for the Class 1A Corrector Magnets

May 5, 2017

This traveler is intended to cover magnetic measurements of the Class 1A corrector magnets.

Receiving:

The following information is to be noted upon receipt of the magnets by the SLAC MFD group:

| Received by (initials): | SDA |
| --- | --- |
| Date placed on test stand (dd-mmm-yyyy): | 11/9/2017 |
| SLAC barcode number: | 4605 |
| Vendor serial number from magnet label: | 004 |
| SLAC approved electrical safety covers? (Y or N): | Y |
| SLAC approved lifting eyes? (Y or N): | n/a |
| Shipping Damage? (Y or N): | N |
| Vendor tests passed on magnet label? (Y or N): | N |
| SLAC drawing number (enter number): | SA-380-314-26 |

Preparation:

A beam direction arrow, with text “beam direction”, is to be applied to the top and/or connector side of the magnet, preferably by stenciling or rubber stamp, or by sticker supplied by LCLS-II.  Place the beam direction arrow pointing as shown in Fig. 1.

| Beam-direction arrow in place (initials): | SDA |
| --- | --- |

Fiducialization:

No fiducialization needs to be done for these small steering coil magnets.

Magnetic Measurements:

Enter URL of on-line magnetic measurements data (please modify or correct if necessary):

http://www-group.slac.stanford.edu/met/MagMeas/MAGDATA/LCLS-II/Corr/4605/

- Determine connection polarities (with power supply outputting positive current) for both the coils, which produce a “positive” field polarity as shown below:

Figure 1.  All steering coils are defined as “positive” and are powered by bipolar supplies.

- Mark the polarities near both pairs of magnet leads with clear “+” and “” labels as shown above.

| Labels and polarities have been marked (initials): | SDA |
| --- | --- |

- Connect the magnet terminals, in the correct polarity as established above, to a bipolar power supply with maximum current I  6 A.

- Measure the dipole field Bx in the center of the gap with Hall probe from 6 to +6 A in 1-A steps, including zero (13 ‘up’ measurements), and then back down from +6 A to 6 in 1-A steps, including zero (13 ‘down’ measurements).

| Filename & run number of Bx data: | Bhvszdat.ru2 |
| --- | --- |

- Measure the dipole field By in the center of the gap with Hall probe from 6 to +6 A in 1-A steps, including zero (13 ‘up’ measurements), and then back down from +6 A to 6 in 1-A steps, including zero (13 ‘down’ measurements).

| Filename & run number of By data: | Bhvszdat.ru1 |
| --- | --- |

- For one magnet only, measure the length-integrated dipole field, Bxdl, from 6 to +6 A in 1-A steps, including zero (13 ‘up’ measurements), and then back down from +6 A to 6 in 1-A steps, including zero (13 ‘down’ measurements).

| Filename & run number of Bxdl data: | N/A |
| --- | --- |

- For one magnet only, measure the length-integrated dipole field, Bydl, from 6 to +6 A in 1-A steps, including zero (13 ‘up’ measurements), and then back down from +6 A to 6 in 1-A steps, including zero (13 ‘down’ measurements).

| Filename & run number of Bydl data: | N/A |
| --- | --- |

- Measure the inductance and resistance of each of the two coils on this magnet:

| Inductance of XCOR coils (mH): | 0.757 mH |
| --- | --- |
| Resistance of XCOR coils (Ohms): | 0.159 Ohm @ 22.1 oC |
| Inductance of YCOR coils (mH): | 0.597 mH |
| Resistance of YCOR coils (Ohms): | 0.147 Ohm @ 22.1 oC |

- Upon completion of tests, send data link to Mark Woodley who will produce a data analysis file.  Place data analysis file in magnetic measurements data directory

| Magnet data accepted and data analysis file produced | Via email |
| --- | --- |

Enter URL of on-line magnetic measurements analysis data :

http://www-group.slac.stanford.edu/met/MagMeas/MAGDATA/LCLS-II/Corr/4605/Type 1a.pptx