## Slide 1
engineering type: 0.276D14-C
SLAC drawing number: Everson Tesla
SLAC magnet bar-code: 2325
vendor serial number: 002
when measured: 23-Feb-2022 09:30:29
when analyzed: 29-Apr-2022
MAD assignment: BCXCBX23
Itrim = 0 (after bipolar standardize)
BL @ Imain = 0:
after bipolar standardize = +27.25 +/- 0.07 G-m
after degauss = +0.05 +/- 0.02 G-m …
pole-tip field @ magnet center B = -0.38 G

## Slide 2
Use harmonic tolerances from SXRSS bends
Harmonics @ r= 10 mm:
|b1/b0| = 0.023 [0.500] %
|b2/b0| = 0.072 [1.000] %
p0 = 1.0344e-02
p1 = 2.2998e-03
p2 = -7.2285e-04
Harmonics @ r= 10 mm:
|b1/b0| = 0.041 [0.500] %
|b2/b0| = 0.066 [1.000] %
p0 = 2.1501e-02
p1 = 4.1012e-03
p2 = -6.6128e-04
wirevsx.ru1
wirevsx.ru2

## Slide 3
Imain = 0 (after bipolar standardize)
wiredat.ru3
|Imain| < 9 A, Itrim = 0
(after bipolar standardize)
name: BCXCBX23T
# of main coil turns: 234
# of trim coil turns: 69/2
ratio of turns (Nmain/Ntrim): 6.7826
ratio of slopes (measured): 6.4997
IMMO (amps): [-2,2]
BMMO (main coil amps): [-0.3077,+0.3077]

## Slide 4
Imain = 0 (after degauss)
wiredat.ru5
|Imain| < 9 A, Itrim = 0
(after bipolar standardize)