## Slide 1
engineering type: 0.276D14-C
SLAC drawing number: Everson Tesla
SLAC magnet bar-code: 2326
vendor serial number: 003
when measured: 28-Feb-2022 10:48:57
when analyzed: 29-Apr-2022
MAD assignment: BCXCBX24
Itrim = 0 (after bipolar standardize)
BL @ Imain = 0:
after bipolar standardize = -26.97 +/- 0.08 G-m
after degauss = -0.01 +/- 0.04 G-m …
pole-tip field @ magnet center B = -0.4 G

## Slide 2
Use harmonic tolerances from SXRSS bends
Harmonics @ r= 10 mm:
|b1/b0| = 0.018 [0.500] %
|b2/b0| = 0.064 [1.000] %
p0 = 1.2945e-02
p1 = 1.8214e-03
p2 = -6.3768e-04
wirevsx.ru1
wirevsx.ru2
p0 = 1.4260e-02
p1 = 2.3723e-03
p2 = -6.2107e-04
Harmonics @ r= 10 mm:
|b1/b0| = 0.024 [0.500] %
|b2/b0| = 0.062 [1.000] %

## Slide 3
Imain = 0 (after bipolar standardize)
wiredat.ru3
|Imain| < 9 A, Itrim = 0
(after bipolar standardize)
name: BCXCBX24T
# of main coil turns: 234
# of trim coil turns: 69/2
ratio of turns (Nmain/Ntrim): 6.7826
ratio of slopes (measured): 6.6558
IMMO (amps): [-2,2]
BMMO (main coil amps): [-0.3004,+0.3004]

## Slide 4
Imain = 0 (after degauss)
wiredat.ru5
|Imain| < 9 A, Itrim = 0
(after bipolar standardize)